## Sheet1
FW Error Enum (from IInstrument.h) | Error Code | What does this mean? | Does it stop the FW/MW from further actions? i.e. fatal error | Will a user input resolve it? E.g. Load a bottle. | MW passes to SW via | Error Type | Stateful / Transient | Is cleared by FW? | Automatic Recoverable by the module? | SW User Message (Y/N)? | User Recommended action - may be multiple | User message | Service call required? | Service Recommended actions | Notes
HANDLER_EJECTOR_FAILED | 1 | Motion failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: cables from power com board/motor board/motor connected, check the drive shaft , check motors is plugged in to the motor control board
HANDLER_X_AXIS_FAILED | 2 | Motion failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: cables from power com board/motor board/motor connected, check that the gears are still locked onto the motor shafts (coarse and fine drive), check motors are plugged in at rear of gripper board, check opto flags are not damaged
HANDLER_Y_AXIS_FAILED | 3 | Motion failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: Objects stuck beneath the handler carriage, y axis resistance (instrument off), lithium grease rails, belt misalignment, cables from power com board/motor board/motor connected.
HANDLER_Z_AXIS_FAILED | 4 | Motion failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: cables from power com board/motor board/motor connected, check lead screw resistance with instrument off. Check alignment of top cap bushing.
HANDLER_EJECTOR_RECOVERED | 5 | Motion failed but then recovered and the recovery was successful | No | N/A | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | N/A |  | N | N/A
HANDLER_X_AXIS_RECOVERED | 6 | Motion failed but then recovered and the recovery was successful | No | N/A | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | N/A |  | N | N/A
HANDLER_Y_AXIS_RECOVERED | 7 | Motion failed but then recovered and the recovery was successful | No | N/A | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | N/A |  | N | N/A
HANDLER_Z_AXIS_RECOVERED | 8 | Motion failed but then recovered and the recovery was successful | No | N/A | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | N/A |  | N | N/A
HANDLER_UNKNOWN_ERROR | 9 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
HANDLER_NOT_CALIBRATED | 10 | The handler has not been calibrated | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Call service to calibrate the system | The Plate Handler is out of calibration. Please call Bio-Rad support to resolve the issue. | Y | Calibrate the handler
HANDLER_EXTEND_FAILED | 11 | Gripper could not extend | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: cables from power com board/motor board/motor connected, check that the gears are still locked onto the motor shafts (coarse and fine drive), check motors are plugged in at rear of gripper board, check opto flags are not damaged
HANDLER_RETRACT_FAILED | 12 | Gripper could not retract | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: cables from power com board/motor board/motor connected, check that the gears are still locked onto the motor shafts (coarse and fine drive), check motors are plugged in at rear of gripper board, check opto flags are not damaged
HANDLER_LOAD_FAILED | 13 | The Gripper Fingers could not close | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: check that the gears are still locked onto the motor shafts and drive shaft, check motor cable and connection, check opto flags are not damaged, check calibration of handler
HANDLER_UNLOAD_FAILED | 14 | The Gipper fingers could not open | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: check that the gears are still locked onto the motor shafts and drive shaft, check motor cable and connection, check opto flags are not damaged, check calibration of handler
HANDLER_NOT_INITIALIZED | 15 | The handler has not been homed | No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | NA |  | NA | NA
HANDLER_IN_ERROR_STATE | 16 | The handler had a fatal error and cannot move | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
HANDLER_PLATE_MISSING | 17 | Plate not detected when it should have been in gripper | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful / Transient | No |  | N | NA |  | NA | NA | Only transient during initialization, but any time afterwards it will be stateful.
HANDLER_ASKED_TO_MOVE_OUT_OF_RANGE | 18 | Requested handler position was out of range | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Power cycle instrument. Check teach positions, check optos Collect logs and escalate issue.
DG_HEAT_STAKER_ERROR | 19 | The DG heat stake failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable for the heater and the thermocouple are connected properly at the DG local controller and the wires are not broken from movement. Check the connection at heater cartridge.
DG_HEAT_STAKER_MOTION_ERROR | 20 | The DG heat stake did not move to the expected position | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check air tubes to HS head, cable to the HS head valve, other leaks in pressure system. Check opto, loom and flag.
DG_DROPLET_GENERATION_ERROR | 21 | Currently unused | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
DG_HEAD_MOTION_ERROR | 22 | The DG head motion failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check air tubes to DG head, cable to the DG head valve, other leaks in pressure system. Check opto, loom and flag.
DG_STAGE_MOTION_ERROR | 23 | The DG stage motion failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the stage stepper motor cable and connection to board. Visual check of stage plate mechanism for damage. When powered off check stage can run freely on the linear rails.
DG_PRESSURE_TIMEOUT_ERROR | 24 | Pressure not reached in time | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: leaks at tube fittings, ensure all push fittings are homed and threaded fittings are tight, use step set to test pressure/vacuum and to help identify any leaks. Check pump is operating, connected, not leaking etc.
DG_VACUUM_TIMEOUT_ERROR | 25 | The DG vacuum timed out | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: leaks at tube fittings, ensure all push fittings are homed and threaded fittings are tight, use step set to test pressure/vacuum and to help identify any leaks. Check pump is operating, connected, not leaking etc.
DG_NO_PLATE_ERROR | 26 | The DG does not see the plate it is supposed to run | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Clean sensor on board via the hole in the DG stage. Check opto operates, replace board. | The plate detect sensor is not inside the stage. It is in front of the stage. We have not had one fail from being dirty. Typically caused by warped plate not sitting flat, or failed flat flex cable.
DG_NO_CATCH_ERROR | 27 | The DG catch is not engaged | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check DG stage for any visible damage missing parts, check the underside spring mechanism. Move stage on and off activation ramps to see if mechanism is working. Check opto operates, replace board.
DG_NO_RESPONCE_ERROR | 28 | One of DG's proxy nodes did not respond to a command | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
DG_NOT_VALID_CHIP_ERROR | 29 | Requested chip out of range | Yes | Abort plate, power cycle machine to clear error, re-add plate with valid parameters | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
DG_TANK_FILL_ERROR | 30 | Oil tank was not filled | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the following: leaks or blockages at tube fittings, ensure all push fittings are homed and threaded fittings are tight, check reservoir (seal), use step set to test pressure/vacuum and to help identify any leaks, check level sensor on reservoir , check if bottle dock filters and replace if necessary. Check bottle venting. Check bottle select valve. Check DG valves (oil trimming pressure/vacuum selector, oil trimming proportional, Oil bottle pressure) are operating and filter on valve.
DG_EMPTY_WASTE_ERROR | 31 | The DG did not receive a response for requesting the waste | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check connection between DG Fluidic Control and BFC boards? Check DG Fluidic Control and BFC boards for damage? Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
DG_EMPTY_WASTE_DENIED | 32 | The waste was not approved for the DG | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | N/A |  | N | N/A
DG_BUSY | 33 | The DG can not do something because it is already doing something | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
DG_HEATER_ERROR | 34 |  | Yes | No |  | Service Call |  | No |  |  | Let current plates finish and power cycle |  | If persistent | Check or replace thermocouples
DG_UNKNOWN_ERROR | 35 | Currently not being published by the firmware | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A
DR_DISCOVERY_FAILED | 36 | During initialization, DR module failed to detect presence of all needed boards from one or more of the stage, detector and fluidics subsystems. This error may be published for each of the three susbystems. | Yes | No |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to fail initialization. Power cycle the instrument. | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check that cables are connected from relevant boards to DR Master (stage, fluidics, LED detector, motor drivers, piston pumps). Check continuity on cables. Visually inspect relevant boards for damage. Replace relevant boards.
DR_STAGE_HOMING_AFTER_ERROR | 37 | DR stage is being homed due to a DR error and will move to plate exchange position for later removal of failed plate. | No | No |  | Service Call | Stateful | Yes |  | N | N/A |  | N | Remove failed plate. Check DR stage functioning by homing stage through MTS. Check DR functioning by running dummy plate.
DR_STAGE_NOT_CALIBRATED | 38 | During initialization, DR detected that the stage has not been calibrated | Yes | No |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to fail initialization. Power cycle the instrument. | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Run DR stage calibration through MTS.
DR_DATA_RECEIVER_NOT_RUNNING | 39 | Currently not being published by the firmware | N/A | N/A |  | Service Call | Stateful | No |  | N | N/A |  | N | N/A (not being published by FW)
DR_DETECTOR_LED_ERROR | 40 | Detector subsystem was not ready when requested to start measuring new well. | Yes | No |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check LED connections, LED detector, replace LED.
DR_NOT_ALL_BOARDS_REGISTERED | 41 | During initialization, DR module detected partial boards and cannot fully operate for droplet detection. This happens after DR_DISCOVERY_FAILED. | Yes | No |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to fail initialization. Power cycle the instrument. | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check that cables are connected from relevant boards to DR Master (stage, fluidics, LED detector, motor drivers, piston pumps). Check continuity on cables. Visually inspect relevant boards for damage. Replace relevant boards.
DR_CONFIG_INVALID | 42 | During initialization, DR module failed to boot up the boards. | Yes | No |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to fail initialization. Power cycle the instrument. | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check that all necessary cables are connected to DR Master. Visually inspect DR Master and other DR boards for damage. Check continuity on cables. Replace relevant boards.
DR_MASTER_IN_ERROR | 43 | DR module has transitioned to an error state caused by a failure in one of its subsystems. | Yes | No |  | Service Call | Stateful | No |  | N | N/A |  |  | Check that all necessary cables are connected to DR Master. Visually inspect DR Master and other DR boards for damage. Check continuity on cables. Replace DR Master.
DR_WAIT_FOR_READ_TIMEOUT | 44 | DR failed to complete measurement for current well due to a failure in the fluidics subsystem during sample reading. | Yes | No |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connections to/from DR Fluidics board. Check DR Fluidics board for damage. Check fluidic components for damage. Check DR fluidic lines for leaks or blockages, including needle. Check needle to plate alignment. Replace DR Fluidics board. | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
DR_DETECTOR_IN_ERROR | 45 | DR failed to initialize detector subsystem or complete measurement for current well due to a failure in the detector subsystem. | Yes | No |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connections to/from Detector board. Inspect optics components for damage/obstructions/loose components/misalignment. Replace Detector board.
DR_FLUIDICS_IN_ERROR | 46 | DR failed to initialize fluidics subsystem or complete measurement for current well due to a failure in the fluidics subsystem during sample preparation. | Yes | No |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connections to/from DR Fluidics board. Check DR Fluidics board for damage. Check fluidic components for damage. Check DR fluidic lines for leaks or blockages, including needle. Check needle to plate alignment. Replace DR Fluidics board.
DR_PID_NOT_SETTLED_ERROR | 47 | During initialization, Detector LED PID loop failed to settle | Yes | No |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to fail initialization. Power cycle the instrument. | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connections to/from DR Detector LED board. Check Detector LED board for damage. Ensure LED optical fibers are not broken or kinked Ensure LED fibers are properly connected and not pulled out partially (visually compare the ferrule length protruding out from receptacle) Ensure LED Fans are spinning properly Ensure there is no stray light falling on the telescope assembly (e.g. operating with the cover removed) If that fails: Replace suspect LED subassembly If that fails: Replace LED Control PCB If that fails: Replace entire Telescope assy
DR_WASTE_BOTTLE_ERROR | 48 | Error occurred while DR was attempting to flush oil to Waste bottles. Could be caused by hardware failures (valves and pumps, for example) of by lack of consumables, or possibly waste overflow situation. | Yes | Possibly |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Ensure there is enough capacity available in Waste bottles. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please verify that the waste bottles have enough capacity available. Allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check fluidic components for damage (pumps, valves, tubing, fittings, manifolds). Check DR fluidic waste lines for leaks or blockages. Attempt to push to waste through MTS. Check that pumps and valves can each operate through MTS. Empty waste overflow reservoir. Check connection between DR Fluidic Control and BFC boards? Check DR Fluidic Control and BFC boards for damage? Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. Can happen if the waste overflow high sensor is triggered and processing is shut down. This resources with be starved for
DR_DR_BOTTLE_ERROR | 49 | Error occurred while DR was attempting to use oil from DR bottles. Could be caused by hardware failures (valves and pumps, for example) of by lack of consumables. | Yes | Possibly |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please verify that the DR oil bottles have enough capacity available. Allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check fluidic components for damage (Peri pump, DR bottle select valves, tubing, fittings, manifolds). Check DR fluidic oil lines for leaks or blockages. Attempt to pump DR oil through MTS. Check that relevant pumps and valves can each operate through MTS. Check filter at inlet to reservoir. Check level sensor is operating.
DR_UNKNOWN_ERROR | 50 | Currently not being published by the firmware | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A (not being published by FW)
TC_UART_MESSAGE_SENT_INTERRUPT_NOT_RECIEVED | 51 | Confirmation for uart messing being sent was not received | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Local Check cable connection between TC Local and TC Controller boards. Inspect boards for damage. Replace TC Local controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_UART_COMM_TIMEDOUT | 52 | Uart response was not received before timeout | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check power connection to TC Controller board. Check cable connection between TC Local and TC Controller boards. Inspect boards for damage. Replace TC Local controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_UNKNOWN_UART_ERROR | 53 | Currently not being published by the firmware | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A
TC_LINK_ERROR_BAD_CALL | 54 | Command from TC Local Controller to Tc board link status was Bad Call | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_BAD_SEND | 55 | Command from TC Local Controller to Tc board link status was Bad Send | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_BAD_LT_CALL | 56 | Command from TC Local Controller to Tc board link status was Bad LT Call | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_RESP_TIMEOUT | 57 | Command from TC Local Controller to Tc board link status was Response timeout | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_BAD_RESP | 58 | Command from TC Local Controller to Tc board link status was Bad Response | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_SLAVE_STATUS | 59 | Command from TC Local Controller to Tc board link status was Slave Status | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_BAD_RTN_IDX | 60 | Command from TC Local Controller to Tc board link status was Bad Return Index | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_BAD_CMN_CHN | 61 | Command from TC Local Controller to Tc board link status was Bad Command | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_UNEXP_RESP | 62 | Command from TC Local Controller to Tc board link status was Unexpected Response | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_BUSY_ERROR | 63 | Command from TC Local Controller to Tc board link status was Busy Error | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_ATTENTION_TO | 64 | Command from TC Local Controller to Tc board link status was Attention To | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_SLAVE_ABORT | 65 | Command from TC Local Controller to Tc board link status was Slave Abort | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_BAD_RCV | 66 | Command from TC Local Controller to Tc board link status was Bad Receive | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_SLAVE_RCV | 67 | Command from TC Local Controller to Tc board link status was Slave Receive | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_LINK_ERROR_SLAVE_SEND | 68 | Command from TC Local Controller to Tc board link status was Slave Send | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_UNKNOWN_LINK_ERROR | 69 | Currently not being published by the firmware | N/A | N/A |  | Service Call | Stateful | No |  | N | N/A |  | N | N/A
IStatusCallback.InstrumentHadError
TC_RUNSTATE_ERROR_RUNNING | 70 | TC state changed to Running instead of Idle, Complete, or Cancel | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_RUNSTATE_ERROR_PAUSED | 71 | TC state changed to Paused instead of Idle, Complete, Running, or Cancel | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_RUNSTATE_ERROR_HOLD | 72 | TC state changed to Hold instead of Idle, Complete, Running, or Cancel | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_RUNSTATE_ERROR_SELF_TEST | 73 | TC state changed to Self Test instead of Running or Cancel | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
TC_RUNSTATE_ERROR_INCUBATE | 74 | TC state changed to Incubate instead of Running or Cancel | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
TC_RUNSTATE_ERROR_ABORT | 75 | TC state changed to Abort instead of Running or Cancel | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check power connection to TC Controller board. Power cycle TC Controller Inspect boards for damage. Replace TC Controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_RUNSTATE_ERROR_INST_CHECK | 76 | TC state changed to Inst Check instead of Running or Cancel | No | No |  | Warning | Transient | No |  | N | NA |  | NA | NA
TC_DOWNLOAD_FAILED_MESSAGE_LENGTH | 77 | TC protocol was not downloaded because the message length was incorrect | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_DOWNLOAD_FAILED_LINK_STATUS | 78 | TC protocol was not downloaded because link status was not No Error | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_DOWNLOAD_FAILED_FUNC_ID | 79 | TC protocol was not downloaded because function id was incorrect | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on TC Controller Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_DOWNLOAD_FAILED_TIMEOUT | 80 | TC protocol was not downloaded because communication timed out | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_UNKNOWN_DOWNLOAD_FAILED | 81 | Currently not being published by the firmware | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A
LOCAL_CONTROLLER_TO_TC_DOWNLOAD_FAILED | 82 | TC protocol download failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_INSTRUMENT_TO_LOCAL_CONTROLLER_FAILED | 83 | TC protocol from middleware to firmware failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check ethernet cable to Xilinx on DR master from PC. Check CAN connection to TC Local Check power connection to TC Controller board. Replace TC Local controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_OPERATION_NOT_ALLOWED_ERROR | 84 | Stop Run command not allowed at given time | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Error is on either TC Local or TC Controller Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC | Drafted, wating on feedback from FW. See https://digitalbio.atlassian.net/browse/DF-1490
TC_UNKNOWN_ERROR | 85 | Currently not being published by the firmware | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A
TC_LID_ERROR_NO_CHIP_DETECTED | 86 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A
TC_LID_ERROR_CHIP_PLACEMENT | 87 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A
TC_LID_UNKNOWN_ERROR | 88 | Currently not being published by the firmware | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A
TC_LID_AXIS_ERROR_FALTED | 89 | Tc lid motion failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connection from TC lid actuator and encoder to TC Local Control board. With instrument powered down, manually check whether TC lid actuator lead nut can be wound up and down (raising and lowering TC lid). Check operation and positioning of TC lid opto. Replace TC lid actuator.
TC_LID_AXIS_ERROR_ENCODER_COUNT_DOES_NOT_MATCH_CSPIN | 90 | Tc lid not at expected location | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connection from TC lid actuator and encoder to TC Local Control board. Check TC Local Control board for damage (especially on cSPIN related components). Inspect TC lid actuator and encoder for damage/loose components. Replace TC lid actuator and/or encoder.
TC_LID_AXIS_ERROR_CSPIN_COUNT_DOES_NOT_MATCH_REQUESTED | 91 | Tc lid not at expected location | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connection from TC lid actuator and encoder to TC Local Control board. Check TC Local Control board for damage (especially on cSPIN related components). Inspect TC lid actuator and encoder for damage/loose components. Replace TC lid actuator and/or encoder.
TC_LID_AXIS_ERROR_TIMED_OUT | 92 | Tc lid motion timed out | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connection from TC lid actuator and encoder to TC Local Control board. Check power connection to TC Local Control board. Check TC Local Control board for damage (especially on cSPIN related components). Inspect TC lid actuator and encoder for damage/loose components. With instrument powered down, manually check whether TC lid actuator lead nut can be wound up and down (raising and lowering TC lid).
TC_LID_AXIS_UNKNOWN_ERROR | 93 | Currently not being published by the firmware | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A
TC_LID_ADC_BASED_MOTION_ERROR_ABORTED_ADC_CHANGED | 94 | Tc lid motion aborted | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check connections between TC load cell, TC lid actuator/encoder and TC Local Controller. Check load cell sub-assy for loose/damaged components or misalignment. Inspect TC lid actuator and encoder for damage/loose components. Replace TC load cell. Replace TC lid actuator.
TC_LID_ADC_BASED_MOTION_ERROR_KILL_SWITCH_STOPPED_MOTOR | 95 | Tc lid motion stopped by kill switch | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check connections between TC load cell, TC lid actuator/encoder and TC Local Controller. Check load cell sub-assy for loose/damaged components or misalignment. Inspect TC lid actuator and encoder for damage/loose components. Replace TC load cell. Replace TC lid actuator.
TC_LID_ADC_BASED_MOTION_ERROR_NOT_ENOUGH_ADC_RANGE | 96 | Tc lid motion failed | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check connections between TC load cell, TC lid actuator/encoder and TC Local Controller. Check load cell sub-assy for loose/damaged components or misalignment. Inspect TC lid actuator and encoder for damage/loose components. Replace TC load cell. Replace TC lid actuator.
TC_LID_ADC_BASED_MOTION_ERROR_CSPIN_COUNT_DOES_NOT_MATCH_REQUESTED | 97 | Tc lid not at expected location | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connection from TC lid actuator and encoder to TC Local Control board. Check TC Local Control board for damage (especially on cSPIN related components). Inspect TC lid actuator and encoder for damage/loose components. Replace TC lid actuator and/or encoder.
TC_LID_ADC_BASED_MOTION_ERROR_ENCODER_COUNT_DOES_NOT_MATCH_CSPIN_AT_BOTTOM | 98 | Tc lid not at expected location | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connection from TC lid actuator and encoder to TC Local Control board. Check TC Local Control board for damage (especially on cSPIN related components). Inspect TC lid actuator and encoder for damage/loose components. Replace TC lid actuator and/or encoder.
TC_LID_ADC_BASED_MOTION_ERROR_ENCOUDER_COUNT_DOES_NOT_MATCH_CSPIN | 99 | Tc lid not at expected location | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable connection from TC lid actuator and encoder to TC Local Control board. Check TC Local Control board for damage (especially on cSPIN related components). Inspect TC lid actuator and encoder for damage/loose components. Replace TC lid actuator and/or encoder.
TC_LID_ADC_BASED_MOTION_UNKNOWN_ERROR | 100 | Currently not being published by the firmware | N/A | N/A | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | N/A |  | N | N/A
PLATE_RESOURCE_DG_MISSING | 101 | The DG station was not discovered during system connect | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the DG local controller has power (LEDs lit, heartbeat blinking 1Hz), Power and CAN cable connected
PLATE_RESOURCE_DG_IN_ERROR | 102 | The DG station is in an error state | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
PLATE_RESOURCE_TC_MISSING | 103 | The TC station was not discovered during system connect | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the TC Local Controller has power (LEDs lit, heartbeat blinking 1Hz), Power and CAN cable connected
PLATE_RESOURCE_TC_IN_ERROR | 104 | The TC station is in an error state | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
PLATE_RESOURCE_DR_MISSING | 105 | The DR station was not discovered during system connect | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the DR master has power (LEDs lit, heartbeat blinking 1Hz), Power and CAN cable connected
PLATE_RESOURCE_DR_IN_ERROR | 106 | The DR station is in an error state | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
PLATE_RESOURCE_PH_MISSING | 107 | The Plate Handler was not discovered during system connect | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the Y Motion, Z Motion and Gripper boards have power (LEDs lit, heartbeat blinking 1Hz), Power and CAN cables connected
PLATE_RESOURCE_PH_IN_ERROR | 108 | The Plate Handler is in an error state | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
PLATE_RESOURCE_INSUFFICIENT_DG_EG_OIL | 109 | Not enough EvaGreen oil to process plate | No | Load EG oil bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Descoped | Descoped |  | Descoped | N/A
PLATE_RESOURCE_INSUFFICIENT_DG_PROBE_OIL | 110 | Not enough Probes oil to process plate | No | Load PROBE oil bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Add extra DG probes oil (additional bottle or full bottle) | There is not enough Droplet Generation Oil for Probes to run the loaded plates. Please add an extra bottle or one with more oil. | N | N/A
PLATE_RESOURCE_INSUFFICIENT_DR_OIL | 111 | Not enough DR oil to process plate | No | Load DR oil bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Add extra DR probes oil (additional bottle or full bottle) | There is not enough Droplet Reader Oil to run the loaded plates. Please add an extra bottle or one with more oil. | N | N/A
PLATE_RESOURCE_INSUFFICIENT_WASTE | 112 | Too much volume in the waste to process plate | No | Load Waste bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Add extra waste capacity (additional bottle or empty bottle) | There is not enough Waste Oil capacity to run the loaded plates. Please add an extra bottle or one with more capacity. | N | N/A
BOTTLE_MISPLACED_P1 | 113 | Non probes bottle in probes position 1 | No | Remove non probes bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | The bottle in Droplet Generation Oil for Probes position 1 is not recognised as a Droplet Generation Oil for Probes bottle. Please try another bottle. | The bottle in Droplet Generation Oil for Probes position 1 is not recognised as a Droplet Generation Oil for Probes bottle. Please try another bottle. | N | N/A
BOTTLE_MISPLACED_P2 | 114 | Non probes bottle in probes position 2 | No | Remove non probes bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | The bottle in Droplet Generation Oil for Probes position 2 is not recognised as a Droplet Generation Oil for Probes bottle. Please try another bottle. | The bottle in Droplet Generation Oil for Probes position 2 is not recognised as a Droplet Generation Oil for Probes bottle. Please try another bottle. | N | N/A
BOTTLE_MISPLACED_E1 | 115 | Non Evagreen bottle in Evagreen position 1 | No | Remove non EG bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
BOTTLE_MISPLACED_E2 | 116 | Non Evagreen bottle in Evagreen position 2 | No | Remove non EG bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
BOTTLE_MISPLACED_D1 | 117 | Non DR bottle in DR position 1 | No | Remove non DR bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | The bottle in Droplet Reader Oil position 1 is not recognised as a Droplet Reader Oil Bottle. Please try another bottle. | The bottle in Droplet Reader Oil position 1 is not recognised as a Droplet Reader Oil Bottle. Please try another bottle. | N | N/A
BOTTLE_MISPLACED_D2 | 118 | Non DR bottle in DR position 2 | No | Remove non DR bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | The bottle in Droplet Reader Oil position 2 is not recognised as a Droplet Reader Oil Bottle. Please try another bottle. | The bottle in Droplet Reader Oil position 2 is not recognised as a Droplet Reader Oil Bottle. Please try another bottle. | N | N/A
BOTTLE_MISPLACED_W1 | 119 | Non Waste bottle in Waste bottle position 1 | No | Remove non waste bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | The bottle in Waste position 1 is not recognised as a waste bottle. Please try another bottle. | The bottle in Waste position 1 is not recognised as a waste bottle. Please try another bottle. | N | N/A
BOTTLE_MISPLACED_W2 | 120 | Non Waste bottle in Waste bottle position 2 | No | Remove non waste bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | The bottle in Waste position 2 is not recognised as a waste bottle. Please try another bottle. | The bottle in Waste position 2 is not recognised as a waste bottle. Please try another bottle. | N | N/A
RFID_STATION_WRONG_PLATE_ID | 121 | The plate scanned does not match the expected plate for a given slot | No | Remove aborted plate | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | N | NA |  | NA | NA
RFID_STATION_READ_FAILED | 122 | After retries the plates RFID tag could not be read | No | Remove aborted plate | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | N | NA |  | NA | NA
RFID_STATION_WRITE_FAILED | 123 | After retries the plates RFID tag could not be written to | No | Remove aborted plate |  | Assistance Needed | Transient | No
// DrDetectorLedDrive Errors (FW)
ERROR_DR_DETECTOR_LED_DRIVE_V24V_OUT_OF_RANGE | 124 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the power cable is seated correctly Replace LED drive board on side of telescope Potentially a faulty Power and Comms distribution board or instrument 24V power supply.
ERROR_DR_DETECTOR_LED_DRIVE_V5V_OUT_OF_RANGE | 125 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_V3p3_OUT_OF_RANGE | 126 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_V12V_OUT_OF_RANGE | 127 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call Stateful | Transient | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_V5p6_OUT_OF_RANGE | 128 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_V4p65_OUT_OF_RANGE | 129 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_V24_CURRENT_OUT_OF_RANGE | 130 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the LED fans are spinning and cable is seated correctly, reseat connector, replace fan if this doesn't resolve the fan spinning issue. Replace the LEDs Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_LOCAL_TEMP_OUT_OF_RANGE | 131 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the Red and Cyan LED fans are spinning and cable is seated correctly, reseat connector, replace fan if this doesn't resolve the fan spinning issue. Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_REMOTE_TEMP_OUT_OF_RANGE | 132 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the Red and Green LED fans are spinning and cable is seated correctly, reseat connector, replace fan if this doesn't resolve the fan spinning issue. Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_FAN_CURRENT_RED_OUT_OF_RANGE | 133 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the Red LED fan is spinning and cable is seated correctly, reseat connector, replace fan if this doesn't resolve the fan spinning issue. Replace the LED Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_FAN_CURRENT_GREEN_OUT_OF_RANGE | 134 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the Cyan LED fan is spinning and cable is seated correctly, reseat connector, replace fan if this doesn't resolve the fan spinning issue. Replace the LED Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_LED_RED_CURRENT_OUT_OF_RANGE | 135 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the Red LED fan is spinning and cable is seated correctly, reseat connector, replace fan if this doesn't resolve the fan spinning issue. Replace the LED Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_LED_GREEN_CURRENT_OUT_OF_RANGE | 136 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the Cyan LED fan is spinning and cable is seated correctly, reseat connector, replace fan if this doesn't resolve the fan spinning issue. Replace the LED Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_LED_RED_DRIVE_TEMP_OUT_OF_RANGE | 137 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the Red LED fan is spinning and cable is seated correctly, reseat connector, replace fan if this doesn't resolve the fan spinning issue. Replace the LED Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_LED_GREEN_DRIVE_TEMP_OUT_OF_RANGE | 138 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the Cyan LED fan is spinning and cable is seated correctly, reseat connector, replace fan if this doesn't resolve the fan spinning issue. Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_RED_FAN_LOCK_ERROR | 139 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the fan is spinning, cable is connected correctly. Replace fan
ERROR_DR_DETECTOR_LED_DRIVE_CYAN_FAN_LOCK_ERROR | 140 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the fan is spinning, cable is connected correctly. Replace fan
ERROR_DR_DETECTOR_LED_DRIVE_PID_UNSETTLED | 141 |  | No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_DETECTOR_LED_DRIVE_PID_FAILED_TIMEOUT | 142 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | 1) Ensure LED optical fibers are not broken or kinked 2) Ensure LED fibers are properly connected and not pulled out partially (visually compare the ferrule length protruding out from receptacle) 3) Ensure LED Fans are spinning properly 4) Ensure there is no stray light falling on the telescope assembly (e.g. operating with the cover removed) If that fails: 5) Replace suspect LED subassembly If that fails: 6) Replace LED Control PCB If that fails: 7) Replace entire Telescope assy
ERROR_DR_DETECTOR_LED_DRIVE_GREEN_LED_DAC_WARNING_LEVEL_REACHED | 143 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | LED intensity is not high enough. Check optic fibres, replace LED.
ERROR_DR_DETECTOR_LED_DRIVE_RED_LED_DAC_WARNING_LEVEL_REACHED | 144 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | LED intensity is not high enough. Check optic fibres, replace LED.
ERROR_DR_DETECTOR_LED_DRIVE_GREEN_LED_DAC_MAX_VALUE_REACHED | 145 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | N | Power Cycle. Call for help |  | Yes | LED intensity is not high enough. Check optic fibres, replace LED.
ERROR_DR_DETECTOR_LED_DRIVE_RED_LED_DAC_MAX_VALUE_REACHED | 146 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | N | Power Cycle. Call for help |  | Yes | LED intensity is not high enough. Check optic fibres, replace LED.
ERROR_DR_DETECTOR_LED_DRIVE_FAN_FAILURE_ERROR | 147 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check the fan is spinning, cable is connected correctly. Replace fan
ERROR_DR_DETECTOR_LED_DRIVE_CONFIGURATION_ERROR | 148 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_DIGITAL_POT_ERROR | 149 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Replace the LED drive board on side of telescope
ERROR_DR_DETECTOR_LED_DRIVE_UNKNOWN_ERROR | 150 |  | No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | N | NA |  | No | NA
// EntryHotel Errors (FW)
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_P1 | 151 | The RFID tag on Probes oil bottle 1 is not supported | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Generator Probes bottle 1 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_P2 | 152 | The RFID tag on Probes oil bottle 2 is not supported | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Generator Probes bottle 2 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_E1 | 153 | The RFID tag on EvaGreen oil bottle 1 is not supported | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_E2 | 154 | The RFID tag on EvaGreen oil bottle 2 is not supported | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_D1 | 155 | The RFID tag on DR oil bottle 1 is not supported | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Reader bottle 1 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_D2 | 156 | The RFID tag on DR oil bottle 2 is not supported | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Reader bottle 2 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_P1 | 157 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | N | N/A |  | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_P2 | 158 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | N | N/A |  | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_E1 | 159 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_E2 | 160 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_D1 | 161 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | N | N/A |  | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_D2 | 162 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | N | N/A |  | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_P1 | 163 | An I2C error occurred while trying to communicate with the RFID module for Probes 1 | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_P2 | 164 | An I2C error occurred while trying to communicate with the RFID module for Probes 2 | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_E1 | 165 | An I2C error occurred while trying to communicate with the RFID module for EvaGreen 1 | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_E2 | 166 | An I2C error occurred while trying to communicate with the RFID module for EvaGreen 2 | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_D1 | 167 | An I2C error occurred while trying to communicate with the RFID module for DR 1 | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_D2 | 168 | An I2C error occurred while trying to communicate with the RFID module for DR 2 | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_P1 | 169 | Communication to Probes 1 RFID module timed out | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_P2 | 170 | Communication to Probes 2 RFID module timed out | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_E1 | 171 | Communication to EvaGreen 1 RFID module timed out | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_E2 | 172 | Communication to EvaGreen 2 RFID module timed out | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_D1 | 173 | Communication to DR1 RFID module timed out | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_D2 | 174 | Communication to DR 2 RFID module timed out | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | No | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_P1 | 175 | Probes 1 bottle RFID tag has a corrupt Read Only section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Generation Oil for Probes bottle 1 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_P2 | 176 | Probes 2 bottle RFID tag has a corrupt Read Only section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Generation Oil for Probes bottle 2 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_E1 | 177 | EvaGreen 1 bottle RFID tag has a corrupt Read Only section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_E2 | 178 | EvaGreen 2 bottle RFID tag has a corrupt Read Only section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_D1 | 179 | DR1 bottle RFID tag has a corrupt Read Only section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Reader bottle 1 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_D2 | 180 | DR 2 bottle RFID tag has a corrupt Read Only section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Reader bottle 2 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_P1 | 181 | Probes 1 bottle RFID tag has a corrupt Read/Write section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Generation Oil for Probes bottle 1 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_P2 | 182 | Probes 2 bottle RFID tag has a corrupt Read/Write section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Generation Oil for Probes bottle 2 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_E1 | 183 | EvaGreen 1 bottle RFID tag has a corrupt Read/Write section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_E2 | 184 | EvaGreen 2 bottle RFID tag has a corrupt Read/Write section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_D1 | 185 | DR1 bottle RFID tag has a corrupt Read/Write section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Reader bottle 1 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_D2 | 186 | DR 2 bottle RFID tag has a corrupt Read/Write section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Droplet Reader bottle 2 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_P1 | 187 | RFID Communication error for specified module. Cause unknown. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_P2 | 188 | RFID Communication error for specified module. Cause unknown. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_E1 | 189 | RFID Communication error for specified module. Cause unknown. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_E2 | 190 | RFID Communication error for specified module. Cause unknown. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_D1 | 191 | RFID Communication error for specified module. Cause unknown. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_D2 | 192 | RFID Communication error for specified module. Cause unknown. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_P1 | 193 | Probes 1 bottle UUID did not match the expected UUID | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_P2 | 194 | Probes 2 bottle UUID did not match the expected UUID | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_E1 | 195 | EvaGreen 1 bottle UUID did not match the expected UUID | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_E2 | 196 | EvaGreen 2 bottle UUID did not match the expected UUID | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_D1 | 197 | DR 1 bottle UUID did not match the expected UUID | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_D2 | 198 | DR 2 bottle UUID did not match the expected UUID | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_P1 | 199 | Probes 1 bottle Read/Write data length exceeded maximum allowed length or tag is not valid | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_P2 | 200 | Probes 2 bottle Read/Write data length exceeded maximum allowed length or tag is not valid | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_E1 | 201 | Eva Green 1 bottle Read/Write data length exceeded maximum allowed length or tag is not valid | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_E2 | 202 | EvaGreen 2 bottle Read/Write data length exceeded maximum allowed length or tag is not valid | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_D1 | 203 | DR 1 bottle Read/Write data length exceeded maximum allowed length or tag is not valid | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_D2 | 204 | DR 2 bottle Read/Write data length exceeded maximum allowed length or tag is not valid | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_P1 | 205 | Probes 1 bottle tag write buffer is corrupt | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_P2 | 206 | Probes 2 bottle tag write buffer is corrupt | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_E1 | 207 | Evagreen 1 bottle tag write buffer is corrupt | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_E2 | 208 | Evagreen 2 bottle tag write buffer is corrupt | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_D1 | 209 | DR 1 bottle tag write buffer is corrupt | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_D2 | 210 | DR 2 bottle tag write buffer is corrupt | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_1_BUS_DOWN | 211 | I2C repeater reports bus locked up on specified RFID module | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_2_BUS_DOWN | 212 | I2C repeater reports bus locked up on specified RFID module | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_3_BUS_DOWN | 213 | I2C repeater reports bus locked up on specified RFID module | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_4_BUS_DOWN | 214 | I2C repeater reports bus locked up on specified RFID module | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_5_BUS_DOWN | 215 | I2C repeater reports bus locked up on specified RFID module | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_6_BUS_DOWN | 216 | I2C repeater reports bus locked up on specified RFID module | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_P1 | 217 | Probes 1 bottle RFID module had more than 3 communication errors | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_P2 | 218 | Probes 2 bottle RFID module had more than 3 communication errors | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_E1 | 219 | Eva Green 1 bottle RFID module had more than 3 communication errors | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_E2 | 220 | EvaGreen 2 bottle RFID module had more than 3 communication errors | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_D1 | 221 | DR 1 bottle RFID module had more than 3 communication errors | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_D2 | 222 | DR 2 bottle RFID module had more than 3 communication errors | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_P1 | 223 | Probes 1 bottle I2C bus is not ready or has an error | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_P2 | 224 | Probes 2 bottle I2C bus is not ready or has an error | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_E1 | 225 | EvaGreen 1 bottle I2C bus is not ready or has an error | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_E2 | 226 | EvaGreen 2 bottle I2C bus is not ready or has an error | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_D1 | 227 | DR 1 bottle I2C bus is not ready or has an error | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_D2 | 228 | DR 2 bottle I2C bus is not ready or has an error | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_P1 | 229 | Communication to Probes bottle 1 display led failed | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_P2 | 230 | Communication to Probes bottle 2 display led failed | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_E1 | 231 | Communication to EvaGreen bottle 1 display led failed | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_E2 | 232 | Communication to EvaGreen bottle 2 display led failed | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_D1 | 233 | Communication to DR bottle 1 display led failed | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_D2 | 234 | Communication to DR bottle 2 display led failed | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_P1 | 235 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_P2 | 236 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_E1 | 237 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_E2 | 238 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_D1 | 239 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_D2 | 240 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT1 | 277 | Entry Slot 1 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from entry slot 1 | There is an unexpected plate in Inbox Slot 1. Please remove it. | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT2 | 278 | Entry Slot 2 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from entry slot 2 | There is an unexpected plate in Inbox Slot 2. Please remove it. | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT3 | 279 | Entry Slot 3 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from entry slot 3 | There is an unexpected plate in Inbox Slot 3. Please remove it. | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT4 | 280 | Entry Slot 4 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from entry slot 4 | There is an unexpected plate in Inbox Slot 4. Please remove it. | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT5 | 281 | Entry Slot 5 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from entry slot 5 | There is an unexpected plate in Inbox Slot 5. Please remove it. | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT1 | 282 | Entry Slot 1 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from entry slot 1 | There is an unexpected plate in Inbox Slot 1. Please remove it. | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT2 | 283 | Entry Slot 2 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from entry slot 2 | There is an unexpected plate in Inbox Slot 2. Please remove it. | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT3 | 284 | Entry Slot 3 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from entry slot 3 | There is an unexpected plate in Inbox Slot 3. Please remove it. | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT4 | 285 | Entry Slot 4 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from entry slot 4 | There is an unexpected plate in Inbox Slot 4. Please remove it. | N | N/A
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT5 | 286 | Entry Slot 5 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from entry slot 5 | There is an unexpected plate in Inbox Slot 5. Please remove it. | N | N/A
// ExitHotel Errors (FW)
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BAD_TAG_TYPE_W1 | 287 | The RFID tag on Waste bottle 1 is not supported | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Waste bottle 1 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BAD_TAG_TYPE_W2 | 288 | The RFID tag on Waste bottle 2 is not supported | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Waste bottle 2 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_TAG_W1 | 289 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | N | N/A |  | N | N/A
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_TAG_W2 | 290 | Currently not being published | N/A | N/A | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | N | N/A |  | N | N/A
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BUS_ERROR_W1 | 291 | An I2C error occurred while trying to communicate with the RFID module for Waste 1 | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BUS_ERROR_W2 | 292 | An I2C error occurred while trying to communicate with the RFID module for Waste 2 | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TIMEOUT_W1 | 293 | Communication to Waste 1 RFID module timed out | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | N/A |  | N | N/A
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TIMEOUT_W2 | 294 | Communication to Waste 2 RFID module timed out | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | N/A |  | N | N/A
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_W1 | 295 | Waste 1 bottle RFID tag has a corrupt Read Only section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Waste bottle 1 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_W2 | 296 | Waste 2 bottle RFID tag has a corrupt Read Only section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Waste bottle 2 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_W1 | 297 | Waste 1 bottle RFID tag has a corrupt Read/Write section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Waste bottle 1 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_W2 | 298 | Waste 2 bottle RFID tag has a corrupt Read/Write section | No | Replace bottle with new bottle | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove indicated bottle and try another one. | The information on Waste bottle 2 could not be read successfully. Please remove it and try another bottle. | If persistent | Return bottle for investigation
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_UNKNOWN_ERROR_W1 | 299 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_UNKNOWN_ERROR_W2 | 300 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_W1 | 301 | Waste 1 bottle UUID did not match the expected UUID | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_W2 | 302 | Waste 2 bottle UUID did not match the expected UUID | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_W1 | 303 | Waste 1 bottle Read/Write data length exceeded maximum allowed length or tag is not valid | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_W2 | 304 | Waste 2 bottle Read/Write data length exceeded maximum allowed length or tag is not valid | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_W1 | 305 | Waste 1 bottle tag write buffer is corrupt | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_W2 | 306 | Waste 2 bottle tag write buffer is corrupt | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_MODULE_1_BUS_DOWN | 307 | I2C repeater reports bus locked up for specified RFID module | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_MODULE_2_BUS_DOWN | 308 | I2C repeater reports bus locked up for specified RFID module | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_W1 | 309 | Waste 1 bottle RFID module had more than 3 communication errors | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_W2 | 310 | Waste 2 bottle RFID module had more than 3 communication errors | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_W1 | 311 | Waste 1 bottle I2C bus is not ready or has an error | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_W2 | 312 | Waste 2 bottle I2C bus is not ready or has an error | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_W1 | 313 | Communication to Waste bottle 1 display led failed | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_W2 | 314 | Communication to Waste bottle 2 display led failed | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_IS_HUNG_W1 | 315 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_IS_HUNG_W2 | 316 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT1 | 329 | Exit Slot 1 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from exit slot 1 | There is an unexpected plate in Outbox Slot 1. Please remove it. | N | N/A
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT2 | 330 | Exit Slot 2 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from exit slot 2 | There is an unexpected plate in Outbox Slot 2. Please remove it. | N | N/A
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT3 | 331 | Exit Slot 3 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from exit slot 3 | There is an unexpected plate in Outbox Slot 3. Please remove it. | N | N/A
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT4 | 332 | Exit Slot 4 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from exit slot 4 | There is an unexpected plate in Outbox Slot 4. Please remove it. | N | N/A
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT5 | 333 | Exit Slot 5 was not empty while trying to reserve it | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | No |  | Y | Remove plate from exit slot 5 | There is an unexpected plate in Outbox Slot 5. Please remove it. | N | N/A
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT1 | 334 | Exit Slot 1 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from exit slot 1 | There is an unexpected plate in Outbox Slot 1. Please remove it. | N | N/A
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT2 | 335 | Exit Slot 2 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from exit slot 2 | There is an unexpected plate in Outbox Slot 2. Please remove it. | N | N/A
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT3 | 336 | Exit Slot 3 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from exit slot 3 | There is an unexpected plate in Outbox Slot 3. Please remove it. | N | N/A
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT4 | 337 | Exit Slot 4 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from exit slot 4 | There is an unexpected plate in Outbox Slot 4. Please remove it. | N | N/A
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT5 | 338 | Exit Slot 5 contains an unexpected plate | No | Yes. Remove plate. | IStatusCallback.InstrumentHadError | Assistance Needed | Stateful | Yes |  | Y | Remove plate from exit slot 5 | There is an unexpected plate in Outbox Slot 5. Please remove it. | N | N/A
// BottleAndFanController Errors (FW)
ERROR_CHASIS_BFC_RFID_ERROR_BAD_TAG_TYPE | 339 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_CORRUPT_TAG | 340 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_BUS_ERROR | 341 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_TIMEOUT | 342 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_CORRUPT_RO_SECTION | 343 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_CORRUPT_RW_SECTION | 344 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_UNKNOWN_ERROR | 359 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE | 346 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_WRITE_DATA_FAILED | 347 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT | 348 | Internal Plate RFID read error? | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_P1 | 360 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | An error has occurred when trying to select a Droplet Generation Oil for Probes bottle. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable to valve. Replace valve.
ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_P2 | 361 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | An error has occurred when trying to select a Droplet Generation Oil for Probes bottle. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable to valve. Replace valve.
ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_E1 | 362 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_E2 | 363 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_P1 | 365 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | An error has occurred when trying to select a Droplet Generation Oil for Probes bottle. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable to valve. Replace valve.
ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_P2 | 366 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | An error has occurred when trying to select a Droplet Generation Oil for Probes bottle. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check cable to valve. Replace valve.
ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_E1 | 367 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_E2 | 368 |  | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_P1 | 370 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_P2 | 371 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_E1 | 372 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_E2 | 373 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_D1 | 374 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_D2 | 375 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_P1 | 382 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_P2 | 383 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_E1 | 384 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_E2 | 385 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_D1 | 386 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_D2 | 387 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_P1 | 388 | BFC command to EntryHotel to write Probes bottle 1 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_P2 | 389 | BFC command to EntryHotel to write Probes bottle 2 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_E1 | 390 | BFC command to EntryHotel to write EvaGreen bottle 1 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_E2 | 391 | BFC command to EntryHotel to write EvaGreen bottle 2 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_D1 | 392 | BFC command to EntryHotel to write DR bottle 1 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_D2 | 393 | BFC command to EntryHotel to write DR bottle 2 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_READ_FAILED_P1 | 394 | BFC command to EntryHotel to read Probes bottle 1 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_READ_FAILED_P2 | 395 | BFC command to EntryHotel to read Probes bottle 2 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_READ_FAILED_E1 | 396 | BFC command to EntryHotel to read EvaGreen bottle 1 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_RFID_READ_FAILED_E2 | 397 | BFC command to EntryHotel to read EvaGreen bottle 2 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | Descoped | Descoped |  | Descoped | N/A
ERROR_CHASIS_BFC_RFID_READ_FAILED_D1 | 398 | BFC command to EntryHotel to read DR bottle 1 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_RFID_READ_FAILED_D2 | 399 | BFC command to EntryHotel to read DR bottle 2 failed. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_DG_WASTE_OVER_TIME_RELEASED_W1 | 402 | No one released the waste within allowed time. Thus BFC released it itself. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_DG_WASTE_OVER_TIME_RELEASED_W2 | 403 | No one released the waste within allowed time. Thus BFC released it itself. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_DR_WASTE_OVER_TIME_RELEASED_W1 | 404 | No one released the waste within allowed time. Thus BFC released it itself. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_DR_WASTE_OVER_TIME_RELEASED_W2 | 405 | No one released the waste within allowed time. Thus BFC released it itself. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_WASTE_OVERFLOW_W1 | 406 |  | Yes | Yes. Empty waste bucket. | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | Notify the user that an overflow has occurred when pumping to waste bottle 1. Extract overflowed waste manually. Replace waste bottle 1 if it occurs again. | A waste overflow has occurred when pumping to waste bottle 1. Please manually empty waste overflow reservoir from the waste removal port on the front of the instrument and replace waste bottle 1. If the error re-occurs, please call service for assistance. | If persistent | Check if the waste overflow reservoir is full. Is the waste bottle missing a straw? Is the sensor triggered (orange LED is lit) when it is empty, replace sensor. Is the cable connected correctly?
ERROR_CHASIS_BFC_WASTE_OVERFLOW_W2 | 407 |  | Yes | Yes. Empty waste bucket. | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | Notify the user that an overflow has occurred when pumping to waste bottle 2. Extract overflowed waste manually. Replace waste bottle 2 if it occurs again. | A waste overflow has occurred when pumping to waste bottle 2. Please manually empty waste overflow reservoir from the waste removal port on the front of the instrument and replace waste bottle 2. If the error re-occurs, please call service for assistance. | If persistent | Check if the waste overflow reservoir is full. Is the waste bottle missing a straw? Is the sensor triggered (orange LED is lit) when it is empty, replace sensor. Is the cable connected correctly?
ERROR_CHASIS_BFC_WASTE_IN_ERROR_STATE_W1 | 408 |  | Yes | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_WASTE_IN_ERROR_STATE_W2 | 409 |  | Yes | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_WASTE_OVERFLOW_SENSOR_DETECTED | 410 |  | Yes | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_V24V_OUT_OF_RANGE | 412 | The 24v provided to the BFC board is out of tolerance | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_V5V_OUT_OF_RANGE | 413 | The 5v provided to the BFC board is out of tolerance | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_V3p3V_OUT_OF_RANGE | 414 | The 3.3v provided to the BFC board is out of tolerance | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_TEMP_OUT_OF_RANGE | 415 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BFC_FAN_FAILURE_DG | 416 | The exhaust fan at the DG location has failed. | No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an internal cooling error has occurred which might impact quality of the data. Call service to have the instrument fixed. | An internal cooling error has occurred which might impact quality of the data. Please call Bio-Rad support to resolve the issue. | Y | Check power connection between DG fan and BFC board. Check continuity on fan cables. Attempt to run fan through MTS. Replace fan.
ERROR_CHASIS_BFC_FAN_FAILURE_TC | 417 | Nothing. We do not use the fan at the TC location. | No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an internal cooling error has occurred which might impact quality of the data. Call service to have the instrument fixed. | An internal cooling error has occurred which might impact quality of the data. Please call Bio-Rad support to resolve the issue. | Y | Check power connection between TC fan and BFC board. Check continuity on fan cables. Attempt to run fan through MTS. Replace fan.
ERROR_CHASIS_BFC_FAN_FAILURE_DR | 418 | The exhaust fan at the DR location has failed. | No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an internal cooling error has occurred which might impact quality of the data. Call service to have the instrument fixed. | An internal cooling error has occurred which might impact quality of the data. Please call Bio-Rad support to resolve the issue. | Y | Check power connection between DR fan and BFC board. Check continuity on fan cables. Attempt to run fan through MTS. Replace fan.
// Front Door Controller Errors (FW)
ERROR_CHASIS_FRONT_DOOR_I24V_OUT_OF_RANGE | 419 | Front Door detected over current condition. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | If persistent | Replace motors and front door board. Assuming doors are not just mecanically jammed.
ERROR_CHASIS_FRONT_DOOR_MOTOR_1_FAULT | 420 | Motor driver chip 1 had a fault. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | If combined with door failing | Replace motor and front door board. Assuming doors are not just mecanically jammed.
ERROR_CHASIS_FRONT_DOOR_MOTOR_2_FAULT | 421 | Motor driver chip 2 had a fault. | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | If combined with door failing | Replace motor and front door board. Assuming doors are not just mecanically jammed.
ERROR_CHASIS_FRONT_DOOR_ENTRY_NOT_HOMED | 557 | The system has not been initalized, or front door board asserted. | Yes | No |  | Assistance Needed | Stateful | No |  |  | Power Cycle the system, and re initalize. | Entry DOOR Error. | If persistent | Check cables goint to front door to ensure steady power to board. If repeating failure, replace front door board, and possably motors. | The syetsm can not start without homign the doors. Thus the only way this is likely is if the front door board reboots, either due to an assetr or power issue. In which case it would power up and need to be homed again.
ERROR_CHASIS_FRONT_DOOR_EXIT_NOT_HOMED | 558 | The system has not been initalized, or front door board asserted. | Yes | No |  | Assistance Needed | Stateful | No |  |  | Power Cycle the system, and re initalize. | Exit Door Error | If persistent | Check cables goint to front door to ensure steady power to board. If repeating failure, replace front door board, and possably motors. | The syetsm can not start without homign the doors. Thus the only way this is likely is if the front door board reboots, either due to an assetr or power issue. In which case it would power up and need to be homed again.
ERROR_CHASIS_FRONT_DOOR_ENTRY_DOOR_ERROR_TIMEOUT | 422 |  | No | Yes. Give door a nudge and retry. | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | An inbox door move has failed. Please try again and call service if it persists | An inbox door move has failed. Please try again and call service if it persists | If persistent | Manually check movement of door on its belt (also check that tab at bottom of door slides along its groove). Check for obstructions in the door's path (loose components, cables, ground wires). Check cable connection between door motor and Front Door board. Check door can open/close through MTS. Check alignment of belt, gears, motor and opto. Check functioning of opto. Check that belt is secured on door bracket. Check belt tension. If neither door can open/close, replace Front Door board.
ERROR_CHASIS_FRONT_DOOR_EXIT_DOOR_ERROR_TIMEOUT | 423 |  | No | Yes. Give door a nudge and retry. | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | An outbox door move has failed. Please try again and call service if it persists | An outbox door move has failed. Please try again and call service if it persists. | If persistent | Manually check movement of door on its belt (also check that tab at bottom of door slides along its groove). Check for obstructions in the door's path (loose components, cables, ground wires). Check cable connection between door motor and Front Door board. Check door can open/close through MTS. Check alignment of belt, gears, motor and opto. Check functioning of opto. Check that belt is secured on door bracket. Check belt tension. If neither door can open/close, replace Front Door board.
ERROR_CHASIS_FRONT_DOOR_ENTRY_DOOR_ERROR_OUT_OF_BOUNDS | 424 |  | No | Yes. Give door a nudge and retry. | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | An inbox door move has failed. Please try again and call service if it persists | An inbox door move has failed. Please try again and call service if it persists | If persistent | Manually check movement of door on its belt (also check that tab at bottom of door slides along its groove). Check for obstructions in the door's path (loose components, cables, ground wires). Check cable connection between door motor and Front Door board. Check door can open/close through MTS. Check alignment of belt, gears, motor and opto. Check functioning of opto. Check that belt is secured on door bracket. Check belt tension. If neither door can open/close, replace Front Door board.
ERROR_CHASIS_FRONT_DOOR_EXIT_DOOR_ERROR_OUT_OF_BOUNDS | 425 |  | No | Yes. Give door a nudge and retry. | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | An outbox door move has failed. Please try again and call service if it persists | An outbox door move has failed. Please try again and call service if it persists. | If persistent | Manually check movement of door on its belt (also check that tab at bottom of door slides along its groove). Check for obstructions in the door's path (loose components, cables, ground wires). Check cable connection between door motor and Front Door board. Check door can open/close through MTS. Check alignment of belt, gears, motor and opto. Check functioning of opto. Check that belt is secured on door bracket. Check belt tension. If neither door can open/close, replace Front Door board.
ERROR_CHASIS_FRONT_DOOR_ENTRY_DOOR_ERROR_DID_NOT_REACH_TARGET_POSITION | 426 |  | No | Yes. Give door a nudge and retry. | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | An inbox door move has failed. Please try again and call service if it persists | An inbox door move has failed. Please try again and call service if it persists | If persistent | Manually check movement of door on its belt (also check that tab at bottom of door slides along its groove). Check for obstructions in the door's path (loose components, cables, ground wires). Check cable connection between door motor and Front Door board. Check door can open/close through MTS. Check alignment of belt, gears, motor and opto. Check functioning of opto. Check that belt is secured on door bracket. Check belt tension. If neither door can open/close, replace Front Door board.
ERROR_CHASIS_FRONT_DOOR_EXIT_DOOR_ERROR_DID_NOT_REACH_TARGET_POSITION | 427 |  | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_BAD_TAG_TYPE | 428 | External Plate RFID error during user scan | No | Scan different plate | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | Attempt to rescan. If that doesn't work try another plate | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | If persistent | Scan another plate. Return the plate for investigation
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_CORRUPT_TAG | 429 | External Plate RFID error during user scan | No | Scan different plate | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | Attempt to rescan. If that doesn't work try another plate | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | If persistent | Scan another plate. Return the plate for investigation
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_BUS_ERROR | 430 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TIMEOUT | 431 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_CORRUPT_RO_SECTION | 432 | External Plate RFID error during user scan | No | Scan different plate | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | Attempt to rescan. If that doesn't work try another plate | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | If persistent | Scan another plate. Return the plate for investigation
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_CORRUPT_RW_SECTION | 433 | External Plate RFID error during user scan | No | Scan different plate | IStatusCallback.InstrumentHadError | Assistance Needed | Transient | No |  | Y | Attempt to rescan. If that doesn't work try another plate | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | If persistent | Scan another plate. Return the plate for investigation
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_UNKNOWN_ERROR | 449 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE | 435 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_WRITE_DATA_FAILED | 436 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT | 437 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_RESETTING_MODULE_POWER | 438 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_POWER_FAULT | 439 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_I2C_RFID_BUS_DOWN | 440 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_I2C_FRONT_LED_BUS_DOWN | 441 | External Plate RFID error during user scan | No | No | IStatusCallback.InstrumentHadError | Warning | Transient | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_RESETTING_I2C_PERIPHERAL | 442 | FDC trying to recover RFID module | No | No |  | Warning | Transient | No |  | N | NA | Chassis front door RFID error, resetting I2C peripheral | If persistent | If repeating, need to investigate why RFID module is not working.
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR | 443 | FDC RFID module had an error | No | No |  | Warning | Transient | No |  | N | NA |  | If persistent | If repeating, need to investigate why RFID module is not working.
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG | 444 | UUID from tag too big. Likely corrupt data | No | No |  | Warning | Transient | No |  | N | NA |  | If persistent | If repeating, need to investigate why RFID module is not working.
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT | 445 | Bad command to RFID module | No | No |  | Warning | Transient | No |  | N | NA |  | If persistent | If repeating, need to investigate why RFID module is not working.
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT | 446 | Timeout trying to talk to RFID module | No | No |  | Warning | Transient | No |  | N | NA |  | If persistent | If repeating, need to investigate why RFID module is not working.
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG | 447 | Poorly formed command to RFID module | No | No |  | Warning | Transient | No |  | N | NA |  | If persistent | If repeating, need to investigate why RFID module is not working.
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_MODULE_IS_HALTED | 448 | RFID module halted unexpectedly | No | No |  | Warning | Transient | No |  | N | NA |  | If persistent | If repeating, need to investigate why RFID module is not working.
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_UNKNOWN_ERROR | 449 | RFID module had unknown error | No | No |  | Warning | Transient | No |  | N | NA |  | If persistent | If repeating, need to investigate why RFID module is not working.
ERROR_CHASIS_POWER_CAN_DIST_INTERLOCK_ACTIVE | 450 | Interlock has cut all power to the embedded system. | Yes | No |  | Service Call |  |  |  | Y | Close front panel. OR call service. If due to front panel opened, close and re start plates. |  | Yes | Try to find out why the interlock triggered. Fix root cause. They will have to try to restart their plates after resolving the issue.
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_P1 | 451 | Invalid FW version on Probes RFID Module 1 (Entry Hotel RFID Module 1) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_P2 | 452 | Invalid FW version on Probes RFID Module 2 (Entry Hotel RFID Module 2) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_E1 | 453 | Invalid FW version on Evagreen RFID Module 1 (Entry Hotel RFID Module 3) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_E2 | 454 | Invalid FW version on Evagreen RFID Module 2 (Entry Hotel RFID Module 4) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_D1 | 455 | Invalid FW version on DR Oil RFID Module 1 (Entry Hotel RFID Module 5) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_D2 | 456 | Invalid FW version on DR Oil RFID Module 2 (Entry Hotel RFID Module 6) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_INVALID_FW_VERSION_W1 | 457 | Invalid FW version on Waste Bottle RFID Module 1 (Exit Hotel RFID Module 1) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_RFID_INVALID_FW_VERSION_W2 | 458 | Invalid FW version on Waste Bottle RFID Module 2 (Exit Hotel RFID Module 2) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_RFID_INVALID_FW_VERSION | 459 | Invalid FW version on External Plate RFID Module (Front Door RFID Module) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_RFID_INVALID_FW_VERSION | 460 | Invalid FW version on Internal Plate RFID Module (Bottle and Fan Controller RFID Module) | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_MISSING_NODE_MASTER_PLACEHOLDER | 461 | Master Placeholder was not discovered during connection | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_MISSING_NODE_INSTRUMENT_MASTER | 462 | Instrument Master was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power to the DR master board. Check ethernet cable to Xilinx on DR master from PC. Check Xilinx board is seated correctly on DR master. Reprogram Xilinx board. Replace DR master and Xilinx board.
ERROR_MISSING_NODE_AXIS_X | 463 | DR Axis X was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power and comms connection to Axis X. Re-flash FW? Replace Axis X board.
ERROR_MISSING_NODE_AXIS_Y | 464 | DR Axis Y was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power and comms connection to Axis Y. Re-flash FW? Replace Axis Y board.
ERROR_MISSING_NODE_AXIS_Z | 465 | DR Axis Z was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power and comms connection to Axis Z. Re-flash FW? Replace Axis Z board.
ERROR_MISSING_NODE_PISTON_PUMP_1 | 466 | Piston Pump 1 was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power and comms connection to Piston Pump1. Re-flash FW? Replace Piston Pump 1 board.
ERROR_MISSING_NODE_PISTON_PUMP_2 | 467 | Piston Pump 2 was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power and comms connection to Piston Pump 2. Re-flash FW? Replace Piston Pump 2 board.
ERROR_MISSING_NODE_DETECTOR_LED_DRIVE | 468 | Detector Led Drive was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power and comms connection to Detector. Re-flash FW? Replace Detector board.
ERROR_MISSING_NODE_DR_FLUIDICS_CONTROLLER | 469 | DR Fluidics was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power and comms connection to Fluidics board. Re-flash FW? Replace Fluidics board.
ERROR_MISSING_NODE_BOTTLE_SWITCHING_VALVE_CONTROLLER | 470 | Bottle Switching Valve was not discovered during connection | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_MISSING_NODE_ENTRY_HOTEL | 471 | Entry Hotel was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power and comms connection to Entry Hotel board. Re-flash FW? Replace Entry Hotel board.
ERROR_MISSING_NODE_EXIT_HOTEL | 472 | Exit Hotel was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check power and comms connection to Exit Hotel board. Re-flash FW? Replace Exit Hotel board.
ERROR_MISSING_NODE_POWER_CAN_DIST | 473 | Power and Can Distribution was not discovered during connection | Yes | No | IStatusCallback.InstrumentHadError | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check connection from DR Master to Power and Can Distribution board. Re-flash FW? Replace board.
ERROR_MISSING_NODE_BUS_SENTINEL | 474 | Bus Sentinel was not discovered during connection | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_INSTRUMENT_MASTER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 475 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_AXIS_X_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 476 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_AXIS_Y_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 477 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_AXIS_Z_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 478 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_PISTON_PUMP_1_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 479 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_PISTON_PUMP_2_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 480 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_DETECTOR_LED_DRIVE_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 481 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DG_FLUIDICS_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 482 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_DETECTOR_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 483 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_FLUIDICS_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 484 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DG_LOCAL_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 485 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_TC_LOCAL_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 486 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_HANDLER_X_AXIS_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 487 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_HANDLER_Y_AXIS_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 488 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_HANDLER_Z_AXIS_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 489 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_HANDLER_UNABLE_TO_CHECK_TC_FOR_PLATE | 490 |  | Yes | No |  | Service Call | Stateful | No |  | Y | Power off, open and check manually. Power on. Call service if error persists. | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Check Handler alignment to TC Check for handler axis movement errors TC lid might be stuck down, impeding the handler access.
ERROR_CHASIS_HANDLER_UNABLE_TO_REMOVE_PLATE_FROM_DG | 491 |  | Yes | Remove plate from DG by hand while power is off |  | Service Call | Stateful | No |  | Y | Power off the instrument, open the front door and remove the plate. Then restart the instrument. | A Plate Handler error has occurred while trying to remove a plate from the Droplet Generator. Please turn off the instrument power, open the front panel and remove the plate manually. If the issue persists, please call Bio-Rad support. | If persistent | Check Handler alignment to DG Check for handler axis movement errors, or presence sensor error. DG stage might have a movement error. DG or HS head stuck down? DG Plate retention mechanism stuck. | A plate the is stuck in a sub-module is possibly an indicator of a larger issue, e.g. the Plate Handler is not taught correctly. Therefore while a user action can resolve the issue immediately, a Service Call is still warranted to investigate the root issue.
ERROR_CHASIS_HANDLER_UNABLE_TO_REMOVE_PLATE_FROM_TC | 492 |  | Yes | Remove plate from TC by hand while power is off |  | Service Call | Stateful | No |  | Y | Power off the instrument, open the front door and remove the plate. Then restart the instrument. | A Plate Handler error has occurred while trying to remove a plate from the Thermal Cycler. Please turn off the instrument power, open the front panel and remove the plate manually. If the issue persists, please call Bio-Rad support. | If persistent | Check Handler alignment to TC Check for handler axis movement errors TC lid might be stuck down, impeding the handler access. TC eject mechanism not functioning. | A plate the is stuck in a sub-module is possibly an indicator of a larger issue, e.g. the Plate Handler is not taught correctly. Therefore while a user action can resolve the issue immediately, a Service Call is still warranted to investigate the root issue.
ERROR_CHASIS_HANDLER_UNABLE_TO_REMOVE_PLATE_FROM_DR | 493 |  | Yes | Remove plate from DR by hand while power is off |  | Service Call | Stateful | No |  | Y | Power off the instrument, open the front door and remove the plate. Then restart the instrument. | A Plate Handler error has occurred while trying to remove a plate from the Droplet Reader. Please turn off the instrument power, open the front panel and remove the plate manually. If the issue persists, please call Bio-Rad support. | If persistent | Check Handler alignment to DR Check for handler axis movement errors DR stage movement issue. DR tubes or wiring caught on handler. | A plate the is stuck in a sub-module is possibly an indicator of a larger issue, e.g. the Plate Handler is not taught correctly. Therefore while a user action can resolve the issue immediately, a Service Call is still warranted to investigate the root issue.
ERROR_TC_EJECTOR_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 494 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_ENTRY_HOTEL_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 495 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_MASTER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 496 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_POWER_CAN_DIST_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 497 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_EXIT_HOTEL_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 498 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 499 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_CHASIS_FRONT_DOOR_BOARD_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 500 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
BUS_SENTINAL_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 501 | The listed node did not respond to a periodic GetStatus command | No | No | IStatusCallback.InstrumentHadError | Warning | Stateful | No |  | N | NA |  | NA | NA
ERROR_DR_INSTRUMENT_MASTER_ASSERT_DETECTED | 502 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DR_AXIS_X_ASSERT_DETECTED | 503 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DR_AXIS_Y_ASSERT_DETECTED | 504 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DR_AXIS_Z_ASSERT_DETECTED | 505 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DR_PISTON_PUMP_1_ASSERT_DETECTED | 506 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DR_PISTON_PUMP_2_ASSERT_DETECTED | 507 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DR_DETECTOR_LED_DRIVE_ASSERT_DETECTED | 508 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DG_FLUIDICS_CONTROLLER_ASSERT_DETECTED | 509 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DR_DETECTOR_ASSERT_DETECTED | 510 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DR_FLUIDICS_CONTROLLER_ASSERT_DETECTED | 511 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DG_LOCAL_CONTROLLER_ASSERT_DETECTED | 512 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_TC_LOCAL_CONTROLLER_ASSERT_DETECTED | 513 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_CHASIS_HANDLER_X_AXIS_ASSERT_DETECTED | 514 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_CHASIS_HANDLER_Y_AXIS_ASSERT_DETECTED | 515 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_CHASIS_HANDLER_Z_AXIS_ASSERT_DETECTED | 516 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_TC_EJECTOR_ASSERT_DETECTED | 517 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_CHASIS_ENTRY_HOTEL_ASSERT_DETECTED | 518 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_DR_MASTER_ASSERT_DETECTED | 519 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_CHASIS_POWER_CAN_DIST_ASSERT_DETECTED | 520 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_CHASIS_EXIT_HOTEL_ASSERT_DETECTED | 521 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_ASSERT_DETECTED | 522 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
ERROR_CHASIS_FRONT_DOOR_BOARD_ASSERT_DETECTED | 523 | The listed node returned assert detected as its status from the heartbeat request | Yes/No | No | IStatusCallback.InstrumentHadError | Service Call | Stateful / Transient | No |  | Y | The user should be prompted to contact Bio-Rad service and let them know of the issue, even if no plates were affected or the system still looks fine. | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used. | Only stateful and only stops MW/FW if MW/FW is in a process that uses that particular FW board. Otherwise it is transient and does NOT stop MW/FW.
BUS_SENTINAL_ASSERT_DETECTED | 524 | The listed node returned assert detected as its status from the heartbeat request | No | No | IStatusCallback.InstrumentHadError | Service Call | Transient | No |  | N | NA |  | NA | NA
ERROR_DR_INSTRUMENT_MASTER_BOOTLOADER_RUNNING | 525 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue. |  | If persistent
ERROR_DR_AXIS_X_BOOTLOADER_RUNNING | 526 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DR_AXIS_Y_BOOTLOADER_RUNNING | 527 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DR_AXIS_Z_BOOTLOADER_RUNNING | 528 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DR_PISTON_PUMP_1_BOOTLOADER_RUNNING | 529 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DR_PISTON_PUMP_2_BOOTLOADER_RUNNING | 530 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DR_DETECTOR_LED_DRIVE_BOOTLOADER_RUNNING | 531 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DG_FLUIDICS_CONTROLLER_BOOTLOADER_RUNNING | 532 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DR_DETECTOR_BOOTLOADER_RUNNING | 533 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DR_FLUIDICS_CONTROLLER_BOOTLOADER_RUNNING | 534 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DG_LOCAL_CONTROLLER_BOOTLOADER_RUNNING | 535 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_TC_LOCAL_CONTROLLER_BOOTLOADER_RUNNING | 536 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_CHASIS_HANDLER_X_AXIS_BOOTLOADER_RUNNING | 537 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_CHASIS_HANDLER_Y_AXIS_BOOTLOADER_RUNNING | 538 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_CHASIS_HANDLER_Z_AXIS_BOOTLOADER_RUNNING | 539 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_TC_EJECTOR_BOOTLOADER_RUNNING | 540 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_CHASIS_ENTRY_HOTEL_BOOTLOADER_RUNNING | 541 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_DR_MASTER_BOOTLOADER_RUNNING | 542 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_CHASIS_POWER_CAN_DIST_BOOTLOADER_RUNNING | 543 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_CHASIS_EXIT_HOTEL_BOOTLOADER_RUNNING | 544 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_BOOTLOADER_RUNNING | 545 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
ERROR_CHASIS_FRONT_DOOR_BOARD_BOOTLOADER_RUNNING | 546 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
BUS_SENTINAL_BOOTLOADER_RUNNING | 547 | The listed node is running in the bootloader sector | Yes | Yes/No |  | Service Call | Stateful | No |  |  | The user should power cycle the system. If the problem persist, the user should be prompted to contact Bio-Rad service and let them know of the issue.
HANDLER_LOCATION_CONFIG_NOT_VALID | 548 | A handler location is not configured | Yes | No |  | Service Call |  | No
HANDLER_INVALID_AXIS | 549 | The selected axis does not exist | Yes | No |  | Service Call |  | No
HANDLER_INVALID_LOCATION | 550 | A handler location is invalid | Yes | No |  | Service Call |  | No
HANDLER_SAVE_LOCATION_FAILED | 551 | An unknown location could not be saved in config | Yes | No |  | Service Call |  | No
HANDLER_PLATE_DETECTED_BEFORE_GET_PLATE | 552 | A plate is in the handler when asked to get a new plate | Yes | No |  | Service Call |  | No
HANDLER_BUSY | 553 | The handler is in use and unable to process move request | Yes | No |  | Service Call |  | No
HANDLER_TIMEOUT | 554 | The handler took too long to reach a location | Yes | No |  | Service Call |  | No
HANDLER_PLATE_DETECTED_AFTER_DROP_OFF | 555 | A plate was left in the handler when it should have been dropped off | Yes | No |  | Service Call |  | No
ERROR_MW_ASSERTED | 556 | Qp asserted in the MW | Yes | No |  | Service Call |  | No
ERROR_CHASIS_FRONT_DOOR_ENTRY_NOT_HOMED | 557
ERROR_CHASIS_FRONT_DOOR_EXIT_NOT_HOMED | 558
ERROR_DR_INSTRUMENT_MASTER_CONNECTION_LOST | 559 | Lost ethernet connection to system | yes | no |  | Service Call |  | No
DR_NO_PLATE | 560 | No plate detected on DR stage after retries | yes | no |  | Service Call |  | No
ERROR_VERSIONS_DO_NOT_MATCH | 561 | FW and MW version mismatch | yes | no |  | Service Call |  | No
ERROR_DR_BULK_TRANSFER_PATH_LENGTH | 562 | well data path/file length exceeds 260 char maximum | No/Yes | shorten path to RIP files |  | Assistance Needed |  |  |  | Y | change location of RIP folders to have a shorter path length
ERROR_DR_BULK_TRANSFER_FILE_SIZE | 563 | well data file size is 0 or greater than 8KB | No/Yes | regenerate rip folder |  | Assistance Needed |  |  |  | Y | regenerate RIP files
ERROR_DR_BULK_TRANSFER_FILE_NOT_FOUND | 564 | missing well data file | No/Yes | regenerate rip folder |  | Assistance Needed |  |  |  | Y | regenerate RIP files
ERROR_DR_BULK_TRANSFER_WELL_DATA_FAILED | 565 | bulk transfer of well data failed | No/Yes |  |  | Assistance Needed |  | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. If the issue persists, please call Bio-Rad support. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
ERROR_DR_BULK_TRANSFER_PLATE_DATA_FAILED | 566 | bulk transfer of plate data failed | yes |  |  | Assistance Needed |  | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. If the issue persists, please call Bio-Rad support. | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
ERROR_CHASIS_BFC_WASTE_OVERFLOW_LOWER_SENSOR_DETECTED, | 567 | Lower Waste overflow sensor active | no | yes |  | Assistance Needed |  | yes |  | y | Empty waste overflow tank | Waste detected in the overflow reservoir. Please perform aspirate vent overflow procedure before adding new plates. This procedure can be performed while the instrument is processing plates. If the issue persists, please call Bio-Rad support.
ERROR_PLATE_RECOVERY_DATABASE_CORRUPTED | 568 | The middleware main and backup plate recovery databases are corrupted. A new database will be created | no | no |  | Warning |  | no
ERROR_CHASIS_BFC_WASTE_OVERFLOW_TOP_SENSOR_DETECTED, | 569 | Top Overflow sensor active | yes | no |  | Service Call |  | yes |  | y | Empty waste overflow tank. Restart system.
ERROR_DR_WELL_ACQUISITION_TIMED_OUT | 570 | Internal middleware timer expired waiting for DR to send events | yes | no |  | Service Call |  | No
ERROR_CHASIS_POWER_CAN_DIST_POWER_NOT_GOOD | 571 |  | yes | no |  | Service Call |  | No
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_AC_VOLTAGE_FAIL | 572 |  | yes | no |  | Service Call |  | No
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_OVER_TEMP | 573 |  | yes | no |  | Service Call |  | No
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_24V | 574 |  | yes | no |  | Service Call |  | No
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_ES | 575 |  | yes | no |  | Service Call |  | No
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_TC | 576 |  | yes | no |  | Service Call |  | No
ERROR_MISSING_NODE_HANDLER_X_AXIS | 577 | Node was not discovered during initialization | yes | no |  | Service Call |  | No
ERROR_MISSING_NODE_HANDLER_Y_AXIS | 578 | Node was not discovered during initialization | yes | no |  | Service Call |  | No
ERROR_MISSING_NODE_HANDLER_Z_AXIS | 579 | Node was not discovered during initialization | yes | no |  | Service Call |  | No
ERROR_MISSING_NODE_TC_EJECTOR | 580 | Node was not discovered during initialization | yes | no |  | Service Call |  | No
ERROR_DG_DG_LOCAL_CONTROLLER_HEATER_KILL_SWITCH_ACTIVE | 581 | DG heater kill switch is active | Yes | No |  | Service Call |  | No
ERROR_DG_DG_LOCAL_CONTROLLER_HEATER_TIMEOUT | 582 | DG HeatStaker heater failed | Yes | No |  | Service Call |  | No
ERROR_DG_DG_LOCAL_CONTROLLER_THERMOCOUPLE_ERROR | 583 | DG heater thermocouple measurements failed to converge. | Yes | No |  | Service Call |  | No | y
ERROR_DG_DG_LOCAL_CONTROLLER_THERMOCOUPLE_1_ERROR | 584 | DG heater thermocouple 1 error | Yes | No |  | Service Call |  | No | y
ERROR_DG_DG_LOCAL_CONTROLLER_THERMOCOUPLE_2_ERROR | 585 | DG heater thermocouple 2 error | Yes | No |  | Service Call |  | No | y
ERROR_DG_DG_Master_DG_IN_ERROR_STATE | 586 | DG in error state | Yes | No |  | Service Call |  | No | y
ERROR_CHASIS_HANDLER_MOTION_FAILED | 587 | Plate handler motion timeout (unsuccessful plate move) | Yes | No |  | Service Call |  | No | n
ERROR_TC_TC_LOCAL_CONTROLLER_TC_ERROR | 588 | Currently unused | Yes | No |  | Service Call |  | No | n
UnknownSubmoduleError | 1499 |  | ? | ? |  | Service Call | Stateful | No |  | Y | Notify the user that an error has caused the system to stop. Allow the current runs to finish, if possible. Then power cycle the instrument. | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | If persistent | Collect logs and escalate issue. Power cycle instrument should clear the condition and allow it to be used.
InvalidInstrumentSerialNumbers, | 1501 | During initilization, the SW was unable to get the instrument serial numbers from the instrument or the serial numbers were empty | Yes | No |  | Service Call |  | No |  | Y | Invalid instrument serial numbers.\n Power cycle the instrument and try again.\n If the issue persists, please call Bio-Rad support. | Invalid instrument serial numbers. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
SW_DG_UnknownError, | 1502 | The DG module reported "Error" state but there is no specific error associated with it | Yes | No |  | Service Call |  | No |  | Y |  | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
SW_TC_UnknownError, | 1503 | The TC module reported "Error" state but there is no specific error associated with it | Yes | No |  | Service Call |  | No |  | Y |  | A Thermal cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
SW_DR_UnknownError, | 1504 | The DR module reported "Error" state but there is no specific error associated with it | Yes | No |  | Service Call |  | No |  | Y |  | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
SW_PlateHandler_UnknownError, | 1505 | The plate handler module reported "Error" state but there is no specific error associated with it | Yes | No |  | Service Call |  | No |  | Y |  | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
SW_InboxDoor_UnknownError, | 1506 |  | Yes | No |  | Service Call |  | No |  | Y |  | Power cycle the instrument, push the doors slowly in a couple of inches and power on the instrument and try again. If the issue persists, please call Bio-Rad support.
SW_OutboxDoor_UnknownError, | 1507 |  | Yes | No |  | Service Call |  | No |  | Y |  | Power cycle the instrument, push the doors slowly in a couple of inches and power on the instrument and try again. If the issue persists, please call Bio-Rad support.
SW_Instrument_UnknownError, | 1508 | The instrument reported "Error" state but there is no specific error associated with it | Yes | Yes | Yes | Yes | Yes | Yes |  | Yes |  | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
SW_Instrument_Offline, | 1509 | SW lost communication with the middleware/instrument | Yes | Yes | Yes | Yes | Yes | Yes |  | Yes |  | Instrument connection has been lost.Please power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
SW_InvalidFwVersion | 1510 | FW version is not compatible with the current SW version | Yes | No |  | Service Call | Stateful | No
SW_TimeoutWaitingForResponseFromFW | 1511 | FW did not respond to a SW request within a certain timeout | No | No |  | ? | Transient | No
SoftwareError | 1999 | Unknown software error | Yes | Yes | Yes | Yes | Yes | Yes |  | Yes |  | Software Error, exception caught. Some data might be usable.View the Event log for recent software exceptions and report to Bio-Rad support.
Note on K1: Marked as orange for a draft response. Go green when it has been checked Vs FW inputs in columns C and D. -Deleted user
Note on O1: Marked orange as a draft. Go green when it has been reviewed -Deleted user
Note on A18: At what point does this get generated? does it stop the handler? -Deleted user

## Sheet2
| Error criteria | Error response
1 | Error occurs in FW but FW can continue working without user intervention required. Did not adversly affect the plate running. | Just log it. No user message. No service intervention
2 | Error requires user input to resolve. The the FW will continue working once the user provide the input. | Ask user to provide the input to solve. E.g. change a bottle. Remove a plate. Log the error.
3 | Error prevents the FW continuing to work. A power cycle may clear the issue. | Notify the user that an error has caused the system to stop for reason x. Please power cycle the instrument and call service if the error persists. Log the error.
4 | Error prevents the FW/HW working and will not be cleared by a power cycle | Notify the user that an error has caused the system to stop for reason x. Please call service to resolve. Log the error.

## Temp
Error | What does this mean? | Recommended Service Action
DR_WAIT_FOR_READ_TIMEOUT | DR failed to complete measurement for current well due to a failure in the fluidics subsystem during sample reading. | Check cable connections to/from DR Fluidics board. Check DR Fluidics board for damage. Check fluidic components for damage. Check DR fluidic lines for leaks or blockages, including needle. Check needle to plate alignment. Replace DR Fluidics board.
TC_UART_MESSAGE_SENT_INTERRUPT_NOT_RECIEVED //35 | Confirmation for uart messing being sent was not received | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_UART_COMM_TIMEDOUT | Uart response was not received before timeout | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_BAD_CALL | Command from TC Local Controller to Tc board link status was Bad Call | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_BAD_SEND | Command from TC Local Controller to Tc board link status was Bad Send | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_BAD_LT_CALL //40 | Command from TC Local Controller to Tc board link status was Bad LT Call | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_RESP_TIMEOUT | Command from TC Local Controller to Tc board link status was Response timeout | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_BAD_RESP | Command from TC Local Controller to Tc board link status was Bad Response | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_SLAVE_STATUS | Command from TC Local Controller to Tc board link status was Slave Status | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_BAD_RTN_IDX | Command from TC Local Controller to Tc board link status was Bad Return Index | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_BAD_CMN_CHN //45 | Command from TC Local Controller to Tc board link status was Bad Command | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_UNEXP_RESP | Command from TC Local Controller to Tc board link status was Unexpected Response | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_BUSY_ERROR | Command from TC Local Controller to Tc board link status was Busy Error | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_ATTENTION_TO | Command from TC Local Controller to Tc board link status was Attention To | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_SLAVE_ABORT | Command from TC Local Controller to Tc board link status was Slave Abort | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_BAD_RCV //50 | Command from TC Local Controller to Tc board link status was Bad Receive | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_SLAVE_RCV | Command from TC Local Controller to Tc board link status was Slave Receive | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_LINK_ERROR_SLAVE_SEND | Command from TC Local Controller to Tc board link status was Slave Send | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_RUNSTATE_ERROR_RUNNING | TC state changed to Running instead of Idle, Complete, or Cancel | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_RUNSTATE_ERROR_PAUSED //55 | TC state changed to Paused instead of Idle, Complete, Running, or Cancel | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_RUNSTATE_ERROR_HOLD | TC state changed to Hold instead of Idle, Complete, Running, or Cancel | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_RUNSTATE_ERROR_ABORT | TC state changed to Abort instead of Running or Cancel | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_DOWNLOAD_FAILED_MESSAGE_LENGTH | TC protocol was not downloaded because the message length was incorrect | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_DOWNLOAD_FAILED_LINK_STATUS | TC protocol was not downloaded because link status was not No Error | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_DOWNLOAD_FAILED_FUNC_ID | TC protocol was not downloaded because function id was incorrect | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_DOWNLOAD_FAILED_TIMEOUT | TC protocol was not downloaded because communication timed out | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
LOCAL_CONTROLLER_TO_TC_DOWNLOAD_FAILED | TC protocol download failed | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_INSTRUMENT_TO_LOCAL_CONTROLLER_FAILED | TC protocol from middleware to firmware failed | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC
TC_OPERATION_NOT_ALLOWED_ERROR | Stop Run command not allowed at given time | Check cable connection between TC Local and TC Controller boards. Check power connection to TC Controller board. Inspect boards for damage. Replace TC Local controller board Replace TC

## OriginalCleanedUp
Error Code Name EXCEL | Error Code Name JSON | Error Code | Error Module | Error Module Suggestions |  | Error Type Excel | ErrorType JSON | Error Displays Msg Excel | User message Excel | Error Displays Msg JSON | ErrorDisplayMsg JSON | Error Type Original | Show Error Message Excel
HANDLER_EJECTOR_FAILED | HANDLER_EJECTOR_FAILED | 1 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_X_AXIS_FAILED | HANDLER_X_AXIS_FAILED | 2 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_Y_AXIS_FAILED | HANDLER_Y_AXIS_FAILED | 3 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_Z_AXIS_FAILED | HANDLER_Z_AXIS_FAILED | 4 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_EJECTOR_RECOVERED | HANDLER_EJECTOR_RECOVERED | 5 | PlateHandler | PlateHandler | TRUE | Warning | Warning | FALSE |  | TRUE | Handler ejector recovered | Warning | N
HANDLER_X_AXIS_RECOVERED | HANDLER_X_AXIS_RECOVERED | 6 | PlateHandler | PlateHandler | TRUE | Warning | Warning | FALSE |  | TRUE | Handler x axis recovered | Warning | N
HANDLER_Y_AXIS_RECOVERED | HANDLER_Y_AXIS_RECOVERED | 7 | PlateHandler | PlateHandler | TRUE | Warning | Warning | FALSE |  | TRUE | Handler y axis recovered | Warning | N
HANDLER_Z_AXIS_RECOVERED | HANDLER_Z_AXIS_RECOVERED | 8 | PlateHandler | PlateHandler | TRUE | Warning | Warning | FALSE |  | TRUE | Handler z axis recovered | Warning | N
HANDLER_UNKNOWN_ERROR | HANDLER_UNKNOWN_ERROR | 9 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_NOT_CALIBRATED | HANDLER_NOT_CALIBRATED | 10 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | The Plate Handler is out of calibration. Please call Bio-Rad support to resolve the issue. | TRUE | The Plate Handler is out of calibration. Please call Bio-Rad support to resolve the issue. | Service Call | Y
HANDLER_EXTEND_FAILED | HANDLER_EXTEND_FAILED | 11 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_RETRACT_FAILED | HANDLER_RETRACT_FAILED | 12 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_LOAD_FAILED | HANDLER_LOAD_FAILED | 13 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_UNLOAD_FAILED | HANDLER_UNLOAD_FAILED | 14 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_NOT_INITIALIZED | HANDLER_NOT_INITIALIZED | 15 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | Handler not initialized | Service Call | N
HANDLER_IN_ERROR_STATE | HANDLER_IN_ERROR_STATE | 16 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
HANDLER_PLATE_MISSING | HANDLER_PLATE_MISSING | 17 | PlateHandler | PlateHandler | TRUE | Warning | Warning | FALSE |  | TRUE | Handler plate missing | Warning | N
HANDLER_ASKED_TO_MOVE_OUT_OF_RANGE | HANDLER_ASKED_TO_MOVE_OUT_OF_RANGE | 18 | PlateHandler | PlateHandler | TRUE | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_HEAT_STAKER_ERROR | DG_HEAT_STAKER_ERROR | 19 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_HEAT_STAKER_MOTION_ERROR | DG_HEAT_STAKER_MOTION_ERROR | 20 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_DROPLET_GENERATION_ERROR | DG_DROPLET_GENERATION_ERROR | 21 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_HEAD_MOTION_ERROR | DG_HEAD_MOTION_ERROR | 22 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_STAGE_MOTION_ERROR | DG_STAGE_MOTION_ERROR | 23 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_PRESSURE_TIMEOUT_ERROR | DG_PRESSURE_TIMEOUT_ERROR | 24 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_VACUUM_TIMEOUT_ERROR | DG_VACUUM_TIMEOUT_ERROR | 25 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_NO_PLATE_ERROR | DG_NO_PLATE_ERROR | 26 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_NO_CATCH_ERROR | DG_NO_CATCH_ERROR | 27 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_NO_RESPONCE_ERROR | DG_NO_RESPONCE_ERROR | 28 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_NOT_VALID_CHIP_ERROR | DG_NOT_VALID_CHIP_ERROR | 29 | DG | DG | TRUE | AssistanceNeeded | AssistanceNeeded | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
DG_TANK_FILL_ERROR | DG_TANK_FILL_ERROR | 30 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_EMPTY_WASTE_ERROR | DG_EMPTY_WASTE_ERROR | 31 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_EMPTY_WASTE_DENIED | DG_EMPTY_WASTE_DENIED | 32 | DG | DG | TRUE | Warning | Warning | FALSE |  | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Warning | N
DG_BUSY | DG_BUSY | 33 | DG | DG | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DG_HEATER_ERROR | DG_HEATER_ERROR | 34 | DG | DG | TRUE | ServiceCall | Warning | FALSE |  | TRUE | DG heater error | Service Call
DG_UNKNOWN_ERROR | DG_UNKNOWN_ERROR | 35 | DG | DG | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | N
DR_DISCOVERY_FAILED | DR_DISCOVERY_FAILED | 36 | DR | DR | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_STAGE_HOMING_AFTER_ERROR | DR_STAGE_HOMING_AFTER_ERROR | 37 | DR | DR | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | N
DR_STAGE_NOT_CALIBRATED | DR_STAGE_NOT_CALIBRATED | 38 | DR | DR | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_DATA_RECEIVER_NOT_RUNNING | DR_DATA_RECEIVER_NOT_RUNNING | 39 | DR | DR | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | N
DR_DETECTOR_LED_ERROR | DR_DETECTOR_LED_ERROR | 40 | DR | DR | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_NOT_ALL_BOARDS_REGISTERED | DR_NOT_ALL_BOARDS_REGISTERED | 41 | DR | DR | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_CONFIG_INVALID | DR_CONFIG_INVALID | 42 | DR | DR | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_MASTER_IN_ERROR | DR_MASTER_IN_ERROR | 43 | DR | DR | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | N
DR_WAIT_FOR_READ_TIMEOUT | DR_WAIT_FOR_READ_TIMEOUT | 44 | DR | DR | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_DETECTOR_IN_ERROR | DR_DETECTOR_IN_ERROR | 45 | DR | DR | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_FLUIDICS_IN_ERROR | DR_FLUIDICS_IN_ERROR | 46 | DR | DR | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_PID_NOT_SETTLED_ERROR | DR_PID_NOT_SETTLED_ERROR | 47 | DR | DR | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_WASTE_BOTTLE_ERROR | DR_WASTE_BOTTLE_ERROR | 48 | DR | DR,Bottles | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please verify that the waste bottles have enough capacity available. Allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please verify that the waste bottles have enough capacity available. Allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_DR_BOTTLE_ERROR | DR_DR_BOTTLE_ERROR | 49 | DR | DR,Bottles | TRUE | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please verify that the DR oil bottles have enough capacity available. Allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please verify that the DR oil bottles have enough capacity available. Allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
DR_UNKNOWN_ERROR | DR_UNKNOWN_ERROR | 50 | DR | DR | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | N
TC_UART_MESSAGE_SENT_INTERRUPT_NOT_RECIEVED | TC_UART_MESSAGE_SENT_INTERRUPT_NOT_RECIEVED | 51 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_UART_COMM_TIMEDOUT | TC_UART_COMM_TIMEDOUT | 52 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_UNKNOWN_UART_ERROR | TC_UNKNOWN_UART_ERROR | 53 | TC | TC | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | TC unknown uart error | Service Call | N
TC_LINK_ERROR_BAD_CALL | TC_LINK_ERROR_BAD_CALL | 54 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_BAD_SEND | TC_LINK_ERROR_BAD_SEND | 55 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_BAD_LT_CALL | TC_LINK_ERROR_BAD_LT_CALL | 56 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_RESP_TIMEOUT | TC_LINK_ERROR_RESP_TIMEOUT | 57 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_BAD_RESP | TC_LINK_ERROR_BAD_RESP | 58 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_SLAVE_STATUS | TC_LINK_ERROR_SLAVE_STATUS | 59 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_BAD_RTN_IDX | TC_LINK_ERROR_BAD_RTN_IDX | 60 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_BAD_CMN_CHN | TC_LINK_ERROR_BAD_CMN_CHN | 61 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_UNEXP_RESP | TC_LINK_ERROR_UNEXP_RESP | 62 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_BUSY_ERROR | TC_LINK_ERROR_BUSY_ERROR | 63 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_ATTENTION_TO | TC_LINK_ERROR_ATTENTION_TO | 64 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_SLAVE_ABORT | TC_LINK_ERROR_SLAVE_ABORT | 65 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_BAD_RCV | TC_LINK_ERROR_BAD_RCV | 66 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_SLAVE_RCV | TC_LINK_ERROR_SLAVE_RCV | 67 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LINK_ERROR_SLAVE_SEND | TC_LINK_ERROR_SLAVE_SEND | 68 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_UNKNOWN_LINK_ERROR | TC_UNKNOWN_LINK_ERROR | 69 | TC | TC | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | TC unknown link error | Service Call | N
TC_RUNSTATE_ERROR_RUNNING | TC_RUNSTATE_ERROR_RUNNING | 70 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_RUNSTATE_ERROR_PAUSED | TC_RUNSTATE_ERROR_PAUSED | 71 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_RUNSTATE_ERROR_HOLD | TC_RUNSTATE_ERROR_HOLD | 72 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_RUNSTATE_ERROR_SELF_TEST | TC_RUNSTATE_ERROR_SELF_TEST | 73 | TC | TC | TRUE | Warning | Warning | FALSE |  | TRUE | TC runstate error self test | Warning | N
TC_RUNSTATE_ERROR_INCUBATE | TC_RUNSTATE_ERROR_INCUBATE | 74 | TC | TC | TRUE | Warning | Warning | FALSE |  | TRUE | TC runstate error incubate | Warning | N
TC_RUNSTATE_ERROR_ABORT | TC_RUNSTATE_ERROR_ABORT | 75 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_RUNSTATE_ERROR_INST_CHECK | TC_RUNSTATE_ERROR_INST_CHECK | 76 | TC | TC | TRUE | Warning | Warning | FALSE |  | TRUE | TC runstate error inst check | Warning | N
TC_DOWNLOAD_FAILED_MESSAGE_LENGTH | TC_DOWNLOAD_FAILED_MESSAGE_LENGTH | 77 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_DOWNLOAD_FAILED_LINK_STATUS | TC_DOWNLOAD_FAILED_LINK_STATUS | 78 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_DOWNLOAD_FAILED_FUNC_ID | TC_DOWNLOAD_FAILED_FUNC_ID | 79 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_DOWNLOAD_FAILED_TIMEOUT | TC_DOWNLOAD_FAILED_TIMEOUT | 80 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_UNKNOWN_DOWNLOAD_FAILED | TC_UNKNOWN_DOWNLOAD_FAILED | 81 | TC | TC | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | N
LOCAL_CONTROLLER_TO_TC_DOWNLOAD_FAILED | LOCAL_CONTROLLER_TO_TC_DOWNLOAD_FAILED | 82 | Instrument | TC | FALSE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_INSTRUMENT_TO_LOCAL_CONTROLLER_FAILED | TC_INSTRUMENT_TO_LOCAL_CONTROLLER_FAILED | 83 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_OPERATION_NOT_ALLOWED_ERROR | TC_OPERATION_NOT_ALLOWED_ERROR | 84 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_UNKNOWN_ERROR | TC_UNKNOWN_ERROR | 85 | TC | TC | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | TC unknown error | Service Call | N
TC_LID_ERROR_NO_CHIP_DETECTED | TC_LID_ERROR_NO_CHIP_DETECTED | 86 | TC | TC | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | TC lid error no chip detected | Service Call | N
TC_LID_ERROR_CHIP_PLACEMENT | TC_LID_ERROR_CHIP_PLACEMENT | 87 | TC | TC | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | TC lid error chip placement | Service Call | N
TC_LID_UNKNOWN_ERROR | TC_LID_UNKNOWN_ERROR | 88 | TC | TC | TRUE | ServiceCall | ServiceCall | FALSE |  | TRUE | TC lid unknown error | Service Call | N
TC_LID_AXIS_ERROR_FALTED | TC_LID_AXIS_ERROR_FALTED | 89 | TC | TC | TRUE | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_AXIS_ERROR_ENCODER_COUNT_DOES_NOT_MATCH_CSPIN | TC_LID_AXIS_ERROR_ENCODER_COUNT_DOES_NOT_MATCH_CSPIN | 90 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_AXIS_ERROR_CSPIN_COUNT_DOES_NOT_MATCH_REQUESTED | TC_LID_AXIS_ERROR_CSPIN_COUNT_DOES_NOT_MATCH_REQUESTED | 91 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_AXIS_ERROR_TIMED_OUT | TC_LID_AXIS_ERROR_TIMED_OUT | 92 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_AXIS_UNKNOWN_ERROR | TC_LID_AXIS_UNKNOWN_ERROR | 93 | TC | TC |  | ServiceCall | ServiceCall | FALSE |  | TRUE | TC lid axis unknown error | Service Call | N
TC_LID_ADC_BASED_MOTION_ERROR_ABORTED_ADC_CHANGED | TC_LID_ADC_BASED_MOTION_ERROR_ABORTED_ADC_CHANGED | 94 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_ADC_BASED_MOTION_ERROR_KILL_SWITCH_STOPPED_MOTOR | TC_LID_ADC_BASED_MOTION_ERROR_KILL_SWITCH_STOPPED_MOTOR | 95 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_ADC_BASED_MOTION_ERROR_NOT_ENOUGH_ADC_RANGE | TC_LID_ADC_BASED_MOTION_ERROR_NOT_ENOUGH_ADC_RANGE | 96 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_ADC_BASED_MOTION_ERROR_CSPIN_COUNT_DOES_NOT_MATCH_REQUESTED | TC_LID_ADC_BASED_MOTION_ERROR_CSPIN_COUNT_DOES_NOT_MATCH_REQUESTED | 97 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_ADC_BASED_MOTION_ERROR_ENCODER_COUNT_DOES_NOT_MATCH_CSPIN_AT_BOTTOM | TC_LID_ADC_BASED_MOTION_ERROR_ENCODER_COUNT_DOES_NOT_MATCH_CSPIN_AT_BOTTOM | 98 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_ADC_BASED_MOTION_ERROR_ENCOUDER_COUNT_DOES_NOT_MATCH_CSPIN | TC_LID_ADC_BASED_MOTION_ERROR_ENCOUDER_COUNT_DOES_NOT_MATCH_CSPIN | 99 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
TC_LID_ADC_BASED_MOTION_UNKNOWN_ERROR | TC_LID_ADC_BASED_MOTION_UNKNOWN_ERROR | 100 | TC | TC |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | N
PLATE_RESOURCE_DG_MISSING | PLATE_RESOURCE_DG_MISSING | 101 | Instrument | DG |  | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
PLATE_RESOURCE_DG_IN_ERROR | PLATE_RESOURCE_DG_IN_ERROR | 102 | Instrument | DG |  | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
PLATE_RESOURCE_TC_MISSING | PLATE_RESOURCE_TC_MISSING | 103 | Instrument | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
PLATE_RESOURCE_TC_IN_ERROR | PLATE_RESOURCE_TC_IN_ERROR | 104 | Instrument | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
PLATE_RESOURCE_DR_MISSING | PLATE_RESOURCE_DR_MISSING | 105 | Instrument | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
PLATE_RESOURCE_DR_IN_ERROR | PLATE_RESOURCE_DR_IN_ERROR | 106 | Instrument | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
PLATE_RESOURCE_PH_MISSING | PLATE_RESOURCE_PH_MISSING | 107 | Instrument | PlateHandler |  | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
PLATE_RESOURCE_PH_IN_ERROR | PLATE_RESOURCE_PH_IN_ERROR | 108 | Instrument | PlateHandler |  | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
PLATE_RESOURCE_INSUFFICIENT_DG_EG_OIL | PLATE_RESOURCE_INSUFFICIENT_DG_EG_OIL | 109 | Instrument | DG,Bottles |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Plate resource insufficient DG/EG oil | Assistance Needed | Descoped
PLATE_RESOURCE_INSUFFICIENT_DG_PROBE_OIL | PLATE_RESOURCE_INSUFFICIENT_DG_PROBE_OIL | 110 | Bottles | DG,Bottles |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is not enough Droplet Generation Oil for Probes to run the loaded plates. Please add an extra bottle or one with more oil. | TRUE | There is not enough Droplet Generation Oil for Probes to run the loaded plates. Please add an extra bottle or one with more oil. | Assistance Needed | Y
PLATE_RESOURCE_INSUFFICIENT_DR_OIL | PLATE_RESOURCE_INSUFFICIENT_DR_OIL | 111 | Bottles | DR,Bottles |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is not enough Droplet Reader Oil to run the loaded plates. Please add an extra bottle or one with more oil. | TRUE | There is not enough Droplet Reader Oil to run the loaded plates. Please add an extra bottle or one with more oil. | Assistance Needed | Y
PLATE_RESOURCE_INSUFFICIENT_WASTE | PLATE_RESOURCE_INSUFFICIENT_WASTE | 112 | Bottles | Bottles |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is not enough Waste Oil capacity to run the loaded plates. Please add an extra bottle or one with more capacity. | TRUE | There is not enough Waste Oil capacity to run the loaded plates. Please add an extra bottle or one with more capacity. | Assistance Needed | Y
BOTTLE_MISPLACED_P1 | BOTTLE_MISPLACED_P1 | 113 | Bottles | Bottles |  | AssistanceNeeded | AssistanceNeeded | TRUE | The bottle in Droplet Generation Oil for Probes position 1 is not recognised as a Droplet Generation Oil for Probes bottle. Please try another bottle. | TRUE | The bottle in Droplet Generation Oil for Probes position 1 is not recognized as a Droplet Generation Oil for Probes bottle. Please try another bottle. | Assistance Needed | Y
BOTTLE_MISPLACED_P2 | BOTTLE_MISPLACED_P2 | 114 | Bottles | Bottles |  | AssistanceNeeded | AssistanceNeeded | TRUE | The bottle in Droplet Generation Oil for Probes position 2 is not recognised as a Droplet Generation Oil for Probes bottle. Please try another bottle. | TRUE | The bottle in Droplet Generation Oil for Probes position 2 is not recognized as a Droplet Generation Oil for Probes bottle. Please try another bottle. | Assistance Needed | Y
BOTTLE_MISPLACED_E1 | BOTTLE_MISPLACED_E1 | 115 | Bottles | Bottles |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | The bottle inserted into the DG Oil EvaGreen 1 slot is not recognized as an EvaGreen Oil bottle. • Remove and insert a bottle with DG EvaGreen Oil. | Assistance Needed | Descoped
BOTTLE_MISPLACED_E2 | BOTTLE_MISPLACED_E2 | 116 | Bottles | Bottles |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | The bottle inserted into the DG Oil EvaGreen 2 slot is not recognized as an EvaGreen Oil bottle. • Remove and insert a bottle with DG EvaGreen Oil. | Assistance Needed | Descoped
BOTTLE_MISPLACED_D1 | BOTTLE_MISPLACED_D1 | 117 | Bottles | Bottles |  | AssistanceNeeded | AssistanceNeeded | TRUE | The bottle in Droplet Reader Oil position 1 is not recognised as a Droplet Reader Oil Bottle. Please try another bottle. | TRUE | The bottle in Droplet Reader Oil position 1 is not recognized as a Droplet Reader Oil Bottle. • Please try another bottle. | Assistance Needed | Y
BOTTLE_MISPLACED_D2 | BOTTLE_MISPLACED_D2 | 118 | Bottles | Bottles |  | AssistanceNeeded | AssistanceNeeded | TRUE | The bottle in Droplet Reader Oil position 2 is not recognised as a Droplet Reader Oil Bottle. Please try another bottle. | TRUE | The bottle in Droplet Reader Oil position 2 is not recognized as a Droplet Reader Oil Bottle. • Please try another bottle. | Assistance Needed | Y
BOTTLE_MISPLACED_W1 | BOTTLE_MISPLACED_W1 | 119 | Bottles | Bottles |  | AssistanceNeeded | AssistanceNeeded | TRUE | The bottle in Waste position 1 is not recognised as a waste bottle. Please try another bottle. | TRUE | The bottle in Waste position 1 is not recognized as a waste bottle. • Remove and insert a Waste bottle. | Assistance Needed | Y
BOTTLE_MISPLACED_W2 | BOTTLE_MISPLACED_W2 | 120 | Bottles | Bottles |  | AssistanceNeeded | AssistanceNeeded | TRUE | The bottle in Waste position 2 is not recognised as a waste bottle. Please try another bottle. | TRUE | The bottle in Waste position 2 is not recognized as a waste bottle. • Remove and insert a Waste bottle. | Assistance Needed | Y
RFID_STATION_WRONG_PLATE_ID | RFID_STATION_WRONG_PLATE_ID | 121 | Instrument | Instrument |  | AssistanceNeeded | Warning | FALSE |  | TRUE | Plate from an entry hotel slot has a different RFID than originally scanned. | Assistance Needed | N
RFID_STATION_READ_FAILED | RFID_STATION_READ_FAILED | 122 | Instrument | Instrument |  | AssistanceNeeded | Warning | FALSE |  | TRUE | RFID station read failed | Assistance Needed | N
RFID_STATION_WRITE_FAILED | RFID_STATION_WRITE_FAILED | 123 | Instrument | Instrument |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | RFID station write failed | Assistance Needed
ERROR_DR_DETECTOR_LED_DRIVE_V24V_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_V24V_OUT_OF_RANGE | 124 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_V5V_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_V5V_OUT_OF_RANGE | 125 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_V3p3_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_V3p3_OUT_OF_RANGE | 126 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_V12V_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_V12V_OUT_OF_RANGE | 127 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_V5p6_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_V5p6_OUT_OF_RANGE | 128 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_V4p65_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_V4p65_OUT_OF_RANGE | 129 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_V24_CURRENT_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_V24_CURRENT_OUT_OF_RANGE | 130 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_LOCAL_TEMP_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_LOCAL_TEMP_OUT_OF_RANGE | 131 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_REMOTE_TEMP_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_REMOTE_TEMP_OUT_OF_RANGE | 132 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_FAN_CURRENT_RED_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_FAN_CURRENT_RED_OUT_OF_RANGE | 133 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_FAN_CURRENT_GREEN_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_FAN_CURRENT_GREEN_OUT_OF_RANGE | 134 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_LED_RED_CURRENT_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_LED_RED_CURRENT_OUT_OF_RANGE | 135 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_LED_GREEN_CURRENT_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_LED_GREEN_CURRENT_OUT_OF_RANGE | 136 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_LED_RED_DRIVE_TEMP_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_LED_RED_DRIVE_TEMP_OUT_OF_RANGE | 137 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_LED_GREEN_DRIVE_TEMP_OUT_OF_RANGE | ERROR_DR_DETECTOR_LED_DRIVE_LED_GREEN_DRIVE_TEMP_OUT_OF_RANGE | 138 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_RED_FAN_LOCK_ERROR | ERROR_DR_DETECTOR_LED_DRIVE_RED_FAN_LOCK_ERROR | 139 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_CYAN_FAN_LOCK_ERROR | ERROR_DR_DETECTOR_LED_DRIVE_CYAN_FAN_LOCK_ERROR | 140 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_PID_UNSETTLED | ERROR_DR_DETECTOR_LED_DRIVE_PID_UNSETTLED | 141 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | DR detector LED drive PID unsettled | Service Call | N
ERROR_DR_DETECTOR_LED_DRIVE_PID_FAILED_TIMEOUT | ERROR_DR_DETECTOR_LED_DRIVE_PID_FAILED_TIMEOUT | 142 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_GREEN_LED_DAC_WARNING_LEVEL_REACHED | ERROR_DR_DETECTOR_LED_DRIVE_GREEN_LED_DAC_WARNING_LEVEL_REACHED | 143 | DR | DR |  | Warning | ServiceCall | FALSE |  | TRUE | DR detector LED drive green LED DAC warning level reached | Warning | N
ERROR_DR_DETECTOR_LED_DRIVE_RED_LED_DAC_WARNING_LEVEL_REACHED | ERROR_DR_DETECTOR_LED_DRIVE_RED_LED_DAC_WARNING_LEVEL_REACHED | 144 | DR | DR |  | Warning | ServiceCall | FALSE |  | TRUE | DR detector LED drive red LED DAC warning level reached | Warning | N
ERROR_DR_DETECTOR_LED_DRIVE_GREEN_LED_DAC_MAX_VALUE_REACHED | ERROR_DR_DETECTOR_LED_DRIVE_GREEN_LED_DAC_MAX_VALUE_REACHED | 145 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | DR detector LED drive green LED DAC max value reached | Service Call | N
ERROR_DR_DETECTOR_LED_DRIVE_RED_LED_DAC_MAX_VALUE_REACHED | ERROR_DR_DETECTOR_LED_DRIVE_RED_LED_DAC_MAX_VALUE_REACHED | 146 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | DR detector LED drive red LED DAC max value reached | Service Call | N
ERROR_DR_DETECTOR_LED_DRIVE_FAN_FAILURE_ERROR | ERROR_DR_DETECTOR_LED_DRIVE_FAN_FAILURE_ERROR | 147 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_CONFIGURATION_ERROR | ERROR_DR_DETECTOR_LED_DRIVE_CONFIGURATION_ERROR | 148 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_DIGITAL_POT_ERROR | ERROR_DR_DETECTOR_LED_DRIVE_DIGITAL_POT_ERROR | 149 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_UNKNOWN_ERROR | ERROR_DR_DETECTOR_LED_DRIVE_UNKNOWN_ERROR | 150 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | DR detector LED drive unknown error | Service Call | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_P1 | 151 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Generator Probes bottle 1 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Generation Oil for Probes bottle 1 could not be read successfully. Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_P2 | 152 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Generator Probes bottle 2 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Generation Oil for Probes bottle 2 could not be read successfully. Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_E1 | 153 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis entry hotel RFID error, bad tag type E1 | Assistance Needed | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_E2 | 154 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis entry hotel RFID error, bad tag type E2 | Assistance Needed | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_D1 | 155 | RfidDrOil1 | RfidDrOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Reader bottle 1 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Reader bottle 1 could not be read successfully. Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BAD_TAG_TYPE_D2 | 156 | RfidDrOil2 | RfidDrOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Reader bottle 2 could not be read successfully. Please remove it and try another bottle. | TRUE | "The information on Droplet Reader bottle 2 could not be read successfully. Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_P1 | 157 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis entry hotel RFID error, corrupt tag P1 | Assistance Needed | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_P2 | 158 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis entry hotel RFID error, corrupt tag P2 | Assistance Needed | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_E1 | 159 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis entry hotel RFID error, corrupt tag E1 | Assistance Needed | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_E2 | 160 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis entry hotel RFID error, corrupt tag E2 | Assistance Needed | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_D1 | 161 | RfidDrOil1 | RfidDrOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis entry hotel RFID error, corrupt tag D1 | Assistance Needed | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_TAG_D2 | 162 | RfidDrOil2 | RfidDrOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis entry hotel RFID error, corrupt tag D2 | Assistance Needed | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_P1 | 163 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, bus error P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_P2 | 164 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, bus error P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_E1 | 165 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, bus error E1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_E2 | 166 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, bus error E2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_D1 | 167 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, bus error D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_BUS_ERROR_D2 | 168 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, bus error D2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_P1 | 169 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, timeout P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_P2 | 170 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, timeout P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_E1 | 171 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, timeout E1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_E2 | 172 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, timeout E2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_D1 | 173 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, timeout D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TIMEOUT_D2 | 174 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, timeout D2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_P1 | 175 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Generation Oil for Probes bottle 1 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Generation Oil for Probes bottle 1 could not be read successfully. • Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_P2 | 176 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Generation Oil for Probes bottle 2 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Generation Oil for Probes bottle 2 could not be read successfully. • Please remove it and try another bottle | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_E1 | 177 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | The bottle inserted into the DG Oil EvaGreen 1 slot has a corrupted Read-Only RFID section. • Remove and insert a bottle with DG EvaGreen Oil. | Assistance Needed | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_E2 | 178 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | The bottle inserted into the DG Oil EvaGreen 2 slot has a corrupted Read-Only RFID section. • Remove and insert a bottle with DG EvaGreen Oil. | Assistance Needed | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_D1 | 179 | RfidDrOil1 | RfidDrOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Reader bottle 1 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Reader bottle 1 could not be read successfully. Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_D2 | 180 | RfidDrOil2 | RfidDrOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Reader bottle 2 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Reader bottle 2 could not be read successfully. Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_P1 | 181 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Generation Oil for Probes bottle 1 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Generation Oil for Probes bottle 1 could not be read successfully. • Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_P2 | 182 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Generation Oil for Probes bottle 2 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Generation Oil for Probes bottle 2 could not be read successfully. • Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_E1 | 183 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | The bottle inserted into the DG Oil EvaGreen 1 slot has a corrupted Read-Write RFID section. • Remove and insert a bottle with DG EvaGreen Oil. | Assistance Needed | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_E2 | 184 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | The bottle inserted into the DG Oil EvaGreen 2 slot has a corrupted Read-Write RFID section. • Remove and insert a bottle with DG EvaGreen Oil. | Assistance Needed | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_D1 | 185 | RfidDrOil1 | RfidDrOil1,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Reader bottle 1 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Reader bottle 1 could not be read successfully. • Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_D2 | 186 | RfidDrOil2 | RfidDrOil2,Inbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Droplet Reader bottle 2 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Droplet Reader bottle 2 could not be read successfully. • Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_P1 | 187 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, unknown error P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_P2 | 188 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, unknown error P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_E1 | 189 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, unknown error E1 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_E2 | 190 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, unknown error E2 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_D1 | 191 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, unknown error D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_UNKNOWN_ERROR_D2 | 192 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, unknown error D2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_P1 | 193 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, tag changed before write P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_P2 | 194 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, tag changed before write P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_E1 | 195 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, tag changed before write E1 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_E2 | 196 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, tag changed before write E2 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_D1 | 197 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, tag changed before write D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_D2 | 198 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, tag changed before write D2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_P1 | 199 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data failed P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_P2 | 200 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data failed P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_E1 | 201 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data failed E1 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_E2 | 202 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data failed E2 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_D1 | 203 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data failed D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_D2 | 204 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data failed D2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_P1 | 205 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data buffer corrupt P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_P2 | 206 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data buffer corrupt P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_E1 | 207 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data buffer corrupt E1 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_E2 | 208 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data buffer corrupt E2 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_D1 | 209 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data buffer corrupt D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_D2 | 210 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, write data buffer corrupt D2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_1_BUS_DOWN | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_1_BUS_DOWN | 211 | RfidProbesOil1 | Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, module 1 bus down | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_2_BUS_DOWN | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_2_BUS_DOWN | 212 | RfidProbesOil2 | Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, module 2 bus down | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_3_BUS_DOWN | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_3_BUS_DOWN | 213 | RfidEvagreenOil1 | Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, module 3 bus down | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_4_BUS_DOWN | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_4_BUS_DOWN | 214 | RfidEvagreenOil2 | Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, module 4 bus down | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_5_BUS_DOWN | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_5_BUS_DOWN | 215 | RfidDrOil1 | Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, module 5 bus down | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_6_BUS_DOWN | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_6_BUS_DOWN | 216 | RfidDrOil2 | Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, module 6 bus down | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_P1 | 217 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting module power P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_P2 | 218 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting module power P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_E1 | 219 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting module power E1 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_E2 | 220 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting module power E2 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_D1 | 221 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting module power D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_D2 | 222 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting module power D2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_P1 | 223 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting I2C peripheral P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_P2 | 224 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting I2C peripheral P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_E1 | 225 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting I2C peripheral E1 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_E2 | 226 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting I2C peripheral E2 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_D1 | 227 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting I2C peripheral D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_D2 | 228 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, resetting I2C peripheral D2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_P1 | 229 | RfidProbesOil1 | RfidProbesOil1,Inbox,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, talking to bottle LED's P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_P2 | 230 | RfidProbesOil2 | RfidProbesOil2,Inbox,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, talking to bottle LED's P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_E1 | 231 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, talking to bottle LED's E1 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_E2 | 232 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, talking to bottle LED's E2 | Warning | Descoped
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_D1 | 233 | RfidDrOil1 | RfidDrOil1,Inbox,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, talking to bottle LED's D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_D2 | 234 | RfidDrOil2 | RfidDrOil2,Inbox,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, talking to bottle LED's D2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_P1 | 235 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, P1 is hung | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_P2 | 236 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, P2 is hung | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_E1 | 237 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, E1 is hung | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_E2 | 238 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, E2 is hung | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_D1 | 239 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, D1 is hung | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_IS_HUNG_D2 | 240 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID error, D2 is hung | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT1 | ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT1 | 277 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 1. Please remove it. | TRUE | There is an unexpected plate in Inbox Slot 1. • Please remove it. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT2 | ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT2 | 278 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 2. Please remove it. | TRUE | There is an unexpected plate in Inbox Slot 2. • Please remove it. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT3 | ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT3 | 279 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 3. Please remove it. | TRUE | There is an unexpected plate in Inbox Slot 3. • Please remove it. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT4 | ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT4 | 280 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 4. Please remove it. | TRUE | There is an unexpected plate in Inbox Slot 4. • Please remove it. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT5 | ERROR_CHASIS_ENTRY_HOTEL_SLOT_RESERVE_FAIL_SLOT5 | 281 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 5. Please remove it. | TRUE | There is an unexpected plate in Inbox Slot 5. • Please remove it. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT1 | ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT1 | 282 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 1. Please remove it. | TRUE | Unexpected plate in Inbox Slot 1. Please remove the plate. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT2 | ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT2 | 283 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 2. Please remove it. | TRUE | Unexpected plate in Inbox Slot 2. Please remove the plate. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT3 | ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT3 | 284 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 3. Please remove it. | TRUE | Unexpected plate in Inbox Slot 3. Please remove the plate. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT4 | ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT4 | 285 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 4. Please remove it. | TRUE | Unexpected plate in Inbox Slot 4. Please remove the plate. | Assistance Needed | Y
ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT5 | ERROR_CHASIS_ENTRY_HOTEL_UNEXPECTED_PLATE_SLOT5 | 286 | InboxSlot | InboxSlot,Inbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Inbox Slot 5. Please remove it. | TRUE | Unexpected plate in Inbox Slot 5. Please remove the plate. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BAD_TAG_TYPE_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BAD_TAG_TYPE_W1 | 287 | RfidWaste1 | RfidWaste1,Outbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Waste bottle 1 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Waste bottle 1 could not be read successfully. Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BAD_TAG_TYPE_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BAD_TAG_TYPE_W2 | 288 | RfidWaste2 | RfidWaste2,Outbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Waste bottle 2 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Waste bottle 2 could not be read successfully. Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_TAG_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_TAG_W1 | 289 | RfidWaste1 | RfidWaste1,Outbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis exit hotel RFID error, corrupt tag W1 | Assistance Needed | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_TAG_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_TAG_W2 | 290 | RfidWaste2 | RfidWaste2,Outbox |  | AssistanceNeeded | AssistanceNeeded | FALSE |  | TRUE | Chassis exit hotel RFID error, corrupt tag W2 | Assistance Needed | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BUS_ERROR_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BUS_ERROR_W1 | 291 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, bus error W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BUS_ERROR_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_BUS_ERROR_W2 | 292 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, bus error W2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TIMEOUT_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TIMEOUT_W1 | 293 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, timeout W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TIMEOUT_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TIMEOUT_W2 | 294 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, timeout W2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_W1 | 295 | RfidWaste1 | RfidWaste1,Outbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Waste bottle 1 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Waste bottle 1 could not be read successfully. • Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RO_SECTION_W2 | 296 | RfidWaste2 | RfidWaste2,Outbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Waste bottle 2 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Waste bottle 2 could not be read successfully. • Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_W1 | 297 | RfidWaste1 | RfidWaste1,Outbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Waste bottle 1 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Waste bottle 1 could not be read successfully. • Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_CORRUPT_RW_SECTION_W2 | 298 | RfidWaste2 | RfidWaste2,Outbox |  | AssistanceNeeded | AssistanceNeeded | TRUE | The information on Waste bottle 2 could not be read successfully. Please remove it and try another bottle. | TRUE | The information on Waste bottle 2 could not be read successfully. • Please remove it and try another bottle. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_UNKNOWN_ERROR_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_UNKNOWN_ERROR_W1 | 299 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, unknown error W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_UNKNOWN_ERROR_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_UNKNOWN_ERROR_W2 | 300 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, unknown error W2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_W1 | 301 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, tag changed before write W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE_W2 | 302 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, tag changed before write W2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_W1 | 303 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, write data failed W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_FAILED_W2 | 304 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, write data failed W2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_W1 | 305 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, write data buffer corrupt W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT_W2 | 306 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, write data buffer corrupt W2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_MODULE_1_BUS_DOWN | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_MODULE_1_BUS_DOWN | 307 | RfidWaste1 | Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, module 1 bus down | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_MODULE_2_BUS_DOWN | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_MODULE_2_BUS_DOWN | 308 | RfidWaste2 | Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, module 2 bus down | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_W1 | 309 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, resetting module power W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_MODULE_POWER_W2 | 310 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, resetting module power W2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_W1 | 311 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, resetting I2C peripheral W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_RESETTING_I2C_PERIPHERAL_W2 | 312 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, resetting I2C peripheral W2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_W1 | 313 | RfidWaste1 | RfidWaste1,Outbox,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, talking to bottle LED's W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TALKING_TO_BOTTLE_LEDS_W2 | 314 | RfidWaste2 | RfidWaste2,Outbox,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, talking to bottle LED's W2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_IS_HUNG_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_IS_HUNG_W1 | 315 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, W1 is hung | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_IS_HUNG_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_IS_HUNG_W2 | 316 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel RFID error, W2 is hung | Warning | N
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT1 | ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT1 | 329 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 1. Please remove it. | TRUE | Failed to reserve exit hotel slot 1. • Remove the plate from exit hotel slot 1. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT2 | ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT2 | 330 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 2. Please remove it. | TRUE | Failed to reserve exit hotel slot 2. • Remove the plate from exit hotel slot 2. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT3 | ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT3 | 331 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 3. Please remove it. | TRUE | Failed to reserve exit hotel slot 3. • Remove the plate from exit hotel slot 3. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT4 | ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT4 | 332 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 4. Please remove it. | TRUE | Failed to reserve exit hotel slot 4. • Remove the plate from exit hotel slot 4. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT5 | ERROR_CHASIS_EXIT_HOTEL_SLOT_RESERVE_FAIL_SLOT5 | 333 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 5. Please remove it. | TRUE | Failed to reserve exit hotel slot 5. • Remove the plate from exit hotel slot 5. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT1 | ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT1 | 334 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 1. Please remove it. | TRUE | Unexpected plate in Outbox Slot 1. Please remove the plate. Click on Open Outbox button if the door is closed. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT2 | ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT2 | 335 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 2. Please remove it. | TRUE | Unexpected plate in Outbox Slot 2. Please remove the plate. Click on Open Outbox button if the door is closed. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT3 | ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT3 | 336 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 3. Please remove it. | TRUE | Unexpected plate in Outbox Slot 3. Please remove the plate. Click on Open Outbox button if the door is closed. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT4 | ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT4 | 337 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 4. Please remove it. | TRUE | Unexpected plate in Outbox Slot 4. Please remove the plate. Click on Open Outbox button if the door is closed. | Assistance Needed | Y
ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT5 | ERROR_CHASIS_EXIT_HOTEL_UNEXPECTED_PLATE_SLOT5 | 338 | OutboxSlot | OutboxSlot,Outbox,Slot |  | AssistanceNeeded | AssistanceNeeded | TRUE | There is an unexpected plate in Outbox Slot 5. Please remove it. | TRUE | Unexpected plate in Outbox Slot 5. Please remove the plate. Click on Open Outbox button if the door is closed. | Assistance Needed | Y
ERROR_CHASIS_BFC_RFID_ERROR_BAD_TAG_TYPE | ERROR_CHASIS_BFC_RFID_ERROR_BAD_TAG_TYPE | 339 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, bad tag type | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_CORRUPT_TAG | ERROR_CHASIS_BFC_RFID_ERROR_CORRUPT_TAG | 340 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, corrupt tag | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_BUS_ERROR | ERROR_CHASIS_BFC_RFID_ERROR_BUS_ERROR | 341 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, bus error | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_TIMEOUT | ERROR_CHASIS_BFC_RFID_ERROR_TIMEOUT | 342 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, timeout | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_CORRUPT_RO_SECTION | ERROR_CHASIS_BFC_RFID_ERROR_CORRUPT_RO_SECTION | 343 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, readonly section | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_CORRUPT_RW_SECTION | ERROR_CHASIS_BFC_RFID_ERROR_CORRUPT_RW_SECTION | 344 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, read/write section | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_UNKNOWN_ERROR | ERROR_CHASIS_BFC_RFID_ERROR_UNKNOWN_ERROR | 359 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, unknown error | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE | ERROR_CHASIS_BFC_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE | 346 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, tag changed before write | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_WRITE_DATA_FAILED | ERROR_CHASIS_BFC_RFID_ERROR_WRITE_DATA_FAILED | 347 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, write data failed | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT | ERROR_CHASIS_BFC_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT | 348 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, data buffer corrupt | Warning | N
ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_P1 | ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_P1 | 360 | BottleAndFanController | BottleAndFanController |  | ServiceCall | ServiceCall | TRUE | An error has occurred when trying to select a Droplet Generation Oil for Probes bottle. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Supply Oil error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_P2 | ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_P2 | 361 | BottleAndFanController | BottleAndFanController |  | ServiceCall | ServiceCall | TRUE | An error has occurred when trying to select a Droplet Generation Oil for Probes bottle. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Supply Oil error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_E1 | ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_E1 | 362 | BottleAndFanController | BottleAndFanController |  | ServiceCall | ServiceCall | FALSE |  | TRUE | Chassis bottle and fan controller rotator lost communication E1 | Service Call | Descoped
ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_E2 | ERROR_CHASIS_BFC_ROTATOR_LOST_COMS_E2 | 363 | BottleAndFanController | BottleAndFanController |  | ServiceCall | ServiceCall | FALSE |  | TRUE | Chassis bottle and fan controller rotator lost communication E2 | Service Call | Descoped
ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_P1 | ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_P1 | 365 | BottleAndFanController | BottleAndFanController |  | ServiceCall | ServiceCall | TRUE | An error has occurred when trying to select a Droplet Generation Oil for Probes bottle. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Supply Oil error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_P2 | ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_P2 | 366 | BottleAndFanController | BottleAndFanController |  | ServiceCall | ServiceCall | TRUE | An error has occurred when trying to select a Droplet Generation Oil for Probes bottle. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Supply Oil error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_E1 | ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_E1 | 367 | BottleAndFanController | BottleAndFanController |  | ServiceCall | ServiceCall | FALSE |  | TRUE | Chassis bottle and fan controller rotator timeout E1 | Service Call | Descoped
ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_E2 | ERROR_CHASIS_BFC_ROTATOR_TIMEOUT_E2 | 368 | BottleAndFanController | BottleAndFanController |  | ServiceCall | ServiceCall | FALSE |  | TRUE | Chassis bottle and fan controller rotator timeout E2 | Service Call | Descoped
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_P1 | ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_P1 | 370 | BottleAndFanController | RfidProbesOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write timed out P1 | Warning | N
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_P2 | ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_P2 | 371 | BottleAndFanController | RfidProbesOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write timed out P2 | Warning | N
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_E1 | ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_E1 | 372 | BottleAndFanController | RfidEvagreenOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write timed out E1 | Warning | Descoped
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_E2 | ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_E2 | 373 | BottleAndFanController | RfidEvagreenOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write timed out E2 | Warning | Descoped
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_D1 | ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_D1 | 374 | BottleAndFanController | RfidDrOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write timed out D1 | Warning | N
ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_D2 | ERROR_CHASIS_BFC_RFID_WRITE_TIMEDOUT_D2 | 375 | BottleAndFanController | RfidDrOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write timed out D2 | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_P1 | ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_P1 | 382 | BottleAndFanController | RfidProbesOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, signal received P1 | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_P2 | ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_P2 | 383 | BottleAndFanController | RfidProbesOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, signal received P2 | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_E1 | ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_E1 | 384 | BottleAndFanController | RfidEvagreenOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error,signal received E1 | Warning | Descoped
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_E2 | ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_E2 | 385 | BottleAndFanController | RfidEvagreenOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, signal received E2 | Warning | Descoped
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_D1 | ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_D1 | 386 | BottleAndFanController | RfidDrOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, signal received D1 | Warning | N
ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_D2 | ERROR_CHASIS_BFC_RFID_ERROR_SIGNAL_RECEIVED_D2 | 387 | BottleAndFanController | RfidDrOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID error, signal received D2 | Warning | N
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_P1 | ERROR_CHASIS_BFC_RFID_WRITE_FAILED_P1 | 388 | BottleAndFanController | RfidProbesOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write failed P1 | Warning | N
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_P2 | ERROR_CHASIS_BFC_RFID_WRITE_FAILED_P2 | 389 | BottleAndFanController | RfidProbesOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write failed P2 | Warning | N
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_E1 | ERROR_CHASIS_BFC_RFID_WRITE_FAILED_E1 | 390 | BottleAndFanController | RfidEvagreenOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write failed E1 | Warning | Descoped
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_E2 | ERROR_CHASIS_BFC_RFID_WRITE_FAILED_E2 | 391 | BottleAndFanController | RfidEvagreenOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write failed E2 | Warning | Descoped
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_D1 | ERROR_CHASIS_BFC_RFID_WRITE_FAILED_D1 | 392 | BottleAndFanController | RfidDrOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write failed D1 | Warning | N
ERROR_CHASIS_BFC_RFID_WRITE_FAILED_D2 | ERROR_CHASIS_BFC_RFID_WRITE_FAILED_D2 | 393 | BottleAndFanController | RfidDrOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID write failed D2 | Warning | N
ERROR_CHASIS_BFC_RFID_READ_FAILED_P1 | ERROR_CHASIS_BFC_RFID_READ_FAILED_P1 | 394 | BottleAndFanController | RfidProbesOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID read failed P1 | Warning | N
ERROR_CHASIS_BFC_RFID_READ_FAILED_P2 | ERROR_CHASIS_BFC_RFID_READ_FAILED_P2 | 395 | BottleAndFanController | RfidProbesOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID read failed P2 | Warning | N
ERROR_CHASIS_BFC_RFID_READ_FAILED_E1 | ERROR_CHASIS_BFC_RFID_READ_FAILED_E1 | 396 | BottleAndFanController | RfidEvagreenOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID read failed E1 | Warning | Descoped
ERROR_CHASIS_BFC_RFID_READ_FAILED_E2 | ERROR_CHASIS_BFC_RFID_READ_FAILED_E2 | 397 | BottleAndFanController | RfidEvagreenOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID read failed E2 | Warning | Descoped
ERROR_CHASIS_BFC_RFID_READ_FAILED_D1 | ERROR_CHASIS_BFC_RFID_READ_FAILED_D1 | 398 | BottleAndFanController | RfidDrOil1,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID read failed D1 | Warning | N
ERROR_CHASIS_BFC_RFID_READ_FAILED_D2 | ERROR_CHASIS_BFC_RFID_READ_FAILED_D2 | 399 | BottleAndFanController | RfidDrOil2,BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID read failed D2 | Warning | N
ERROR_CHASIS_BFC_DG_WASTE_OVER_TIME_RELEASED_W1 | ERROR_CHASIS_BFC_DG_WASTE_OVER_TIME_RELEASED_W1 | 402 | BottleAndFanController | BottleAndFanController,DG |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller DG waste over time released W1 | Warning | N
ERROR_CHASIS_BFC_DG_WASTE_OVER_TIME_RELEASED_W2 | ERROR_CHASIS_BFC_DG_WASTE_OVER_TIME_RELEASED_W2 | 403 | BottleAndFanController | BottleAndFanController,DG |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller DG waste over time released W2 | Warning | N
ERROR_CHASIS_BFC_DR_WASTE_OVER_TIME_RELEASED_W1 | ERROR_CHASIS_BFC_DR_WASTE_OVER_TIME_RELEASED_W1 | 404 | BottleAndFanController | BottleAndFanController,DR |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller DR waste over time released W1 | Warning | N
ERROR_CHASIS_BFC_DR_WASTE_OVER_TIME_RELEASED_W2 | ERROR_CHASIS_BFC_DR_WASTE_OVER_TIME_RELEASED_W2 | 405 | BottleAndFanController | BottleAndFanController,DR |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller DR waste over time released W2 | Warning | N
ERROR_CHASIS_BFC_WASTE_OVERFLOW_W1 | ERROR_CHASIS_BFC_WASTE_OVERFLOW_W1 | 406 | BottleAndFanController | BottleAndFanController |  | AssistanceNeeded | AssistanceNeeded | TRUE | A waste overflow has occurred when pumping to waste bottle 1. Please manually empty waste overflow reservoir from the waste removal port on the front of the instrument and replace waste bottle 1. If the error re-occurs, please call service for assistance. | TRUE | A waste overflow has occurred when pumping to waste bottle 1. Please manually empty waste overflow reservoir from the waste removal port on the front of the instrument and replace waste bottle 1. If the error re-occurs, please call service for assistance. | Assistance Needed | Y
ERROR_CHASIS_BFC_WASTE_OVERFLOW_W2 | ERROR_CHASIS_BFC_WASTE_OVERFLOW_W2 | 407 | BottleAndFanController | BottleAndFanController |  | AssistanceNeeded | AssistanceNeeded | TRUE | A waste overflow has occurred when pumping to waste bottle 2. Please manually empty waste overflow reservoir from the waste removal port on the front of the instrument and replace waste bottle 2. If the error re-occurs, please call service for assistance. | TRUE | A waste overflow has occurred when pumping to waste bottle 2. Please manually empty waste overflow reservoir from the waste removal port on the front of the instrument and replace waste bottle 2. If the error re-occurs, please call service for assistance. | Assistance Needed | Y
ERROR_CHASIS_BFC_WASTE_IN_ERROR_STATE_W1 | ERROR_CHASIS_BFC_WASTE_IN_ERROR_STATE_W1 | 408 | BottleAndFanController | BottleAndFanController |  | Warning | AssistanceNeeded | FALSE |  | TRUE | Chassis bottle and fan controller waste in error state W1 | Warning | N
ERROR_CHASIS_BFC_WASTE_IN_ERROR_STATE_W2 | ERROR_CHASIS_BFC_WASTE_IN_ERROR_STATE_W2 | 409 | BottleAndFanController | BottleAndFanController |  | Warning | AssistanceNeeded | FALSE |  | TRUE | Chassis bottle and fan controller waste in error state W2 | Warning | N
ERROR_CHASIS_BFC_WASTE_OVERFLOW_SENSOR_DETECTED | ERROR_CHASIS_BFC_WASTE_OVERFLOW_SENSOR_DETECTED | 410 | BottleAndFanController | BottleAndFanController |  | Warning | AssistanceNeeded | FALSE |  | TRUE | A waste overflow has occurred. Please ensure both waste bottles are empty and manually empty waste reservoir from the waste removal port per the user manual. If the error re-occurs, please call service for assistance. | Warning | N
ERROR_CHASIS_BFC_V24V_OUT_OF_RANGE | ERROR_CHASIS_BFC_V24V_OUT_OF_RANGE | 412 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller V24V out of range | Warning | N
ERROR_CHASIS_BFC_V5V_OUT_OF_RANGE | ERROR_CHASIS_BFC_V5V_OUT_OF_RANGE | 413 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller V5V out of range | Warning | N
ERROR_CHASIS_BFC_V3p3V_OUT_OF_RANGE | ERROR_CHASIS_BFC_V3p3V_OUT_OF_RANGE | 414 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller V3p3V out of range | Warning | N
ERROR_CHASIS_BFC_TEMP_OUT_OF_RANGE | ERROR_CHASIS_BFC_TEMP_OUT_OF_RANGE | 415 | BottleAndFanController | BottleAndFanController |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller temperature out of range | Warning | N
ERROR_CHASIS_BFC_FAN_FAILURE_DG | ERROR_CHASIS_BFC_FAN_FAILURE_DG | 416 | BottleAndFanController | BottleAndFanController,DG |  | ServiceCall | ServiceCall | TRUE | An internal cooling error has occurred which might impact quality of the data. Please call Bio-Rad support to resolve the issue. | TRUE | An internal cooling error has occurred which might impact quality of the data. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_BFC_FAN_FAILURE_TC | ERROR_CHASIS_BFC_FAN_FAILURE_TC | 417 | BottleAndFanController | BottleAndFanController,TC |  | ServiceCall | ServiceCall | TRUE | An internal cooling error has occurred which might impact quality of the data. Please call Bio-Rad support to resolve the issue. | TRUE | An internal cooling error has occurred which might impact quality of the data. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_BFC_FAN_FAILURE_DR | ERROR_CHASIS_BFC_FAN_FAILURE_DR | 418 | BottleAndFanController | BottleAndFanController,DR |  | ServiceCall | ServiceCall | TRUE | An internal cooling error has occurred which might impact quality of the data. Please call Bio-Rad support to resolve the issue. | TRUE | An internal cooling error has occurred which might impact quality of the data. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_FRONT_DOOR_I24V_OUT_OF_RANGE | ERROR_CHASIS_FRONT_DOOR_I24V_OUT_OF_RANGE | 419 | FrontDoor | FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door I24V out of range | Warning | N
ERROR_CHASIS_FRONT_DOOR_MOTOR_1_FAULT | ERROR_CHASIS_FRONT_DOOR_MOTOR_1_FAULT | 420 | InboxDoor | FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door motor 1 fault | Warning | N
ERROR_CHASIS_FRONT_DOOR_MOTOR_2_FAULT | ERROR_CHASIS_FRONT_DOOR_MOTOR_2_FAULT | 421 | OutboxDoor | FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door motor 2 fault | Warning | N
ERROR_CHASIS_FRONT_DOOR_ENTRY_NOT_HOMED | ERROR_CHASIS_FRONT_DOOR_ENTRY_NOT_HOMED | 557 | InboxDoor | FrontDoor |  | AssistanceNeeded | AssistanceNeeded | FALSE | Entry DOOR Error. | TRUE | Entry door motor not homed | Assistance Needed
ERROR_CHASIS_FRONT_DOOR_EXIT_NOT_HOMED | ERROR_CHASIS_FRONT_DOOR_EXIT_NOT_HOMED | 558 | OutboxDoor | FrontDoor |  | AssistanceNeeded | AssistanceNeeded | FALSE | Exit Door Error | TRUE | Exit door motor not homed | Assistance Needed
ERROR_CHASIS_FRONT_DOOR_ENTRY_DOOR_ERROR_TIMEOUT | ERROR_CHASIS_FRONT_DOOR_ENTRY_DOOR_ERROR_TIMEOUT | 422 | InboxDoor | FrontDoor |  | AssistanceNeeded | AssistanceNeeded | TRUE | An inbox door move has failed. Please try again and call service if it persists | TRUE | An inbox door move has failed. Please try again and call service if it persists. | Assistance Needed | Y
ERROR_CHASIS_FRONT_DOOR_EXIT_DOOR_ERROR_TIMEOUT | ERROR_CHASIS_FRONT_DOOR_EXIT_DOOR_ERROR_TIMEOUT | 423 | OutboxDoor | FrontDoor |  | AssistanceNeeded | AssistanceNeeded | TRUE | An outbox door move has failed. Please try again and call service if it persists. | TRUE | An outbox door move has failed. Please try again and call service if it persists. | Assistance Needed | Y
ERROR_CHASIS_FRONT_DOOR_ENTRY_DOOR_ERROR_OUT_OF_BOUNDS | ERROR_CHASIS_FRONT_DOOR_ENTRY_DOOR_ERROR_OUT_OF_BOUNDS | 424 | InboxDoor | FrontDoor |  | AssistanceNeeded | AssistanceNeeded | TRUE | An inbox door move has failed. Please try again and call service if it persists | TRUE | An inbox door move has failed. Please try again and call service if it persists. | Assistance Needed | Y
ERROR_CHASIS_FRONT_DOOR_EXIT_DOOR_ERROR_OUT_OF_BOUNDS | ERROR_CHASIS_FRONT_DOOR_EXIT_DOOR_ERROR_OUT_OF_BOUNDS | 425 | OutboxDoor | FrontDoor |  | AssistanceNeeded | AssistanceNeeded | TRUE | An outbox door move has failed. Please try again and call service if it persists. | TRUE | An outbox door move has failed. Please try again and call service if it persists. | Assistance Needed | Y
ERROR_CHASIS_FRONT_DOOR_ENTRY_DOOR_ERROR_DID_NOT_REACH_TARGET_POSITION | ERROR_CHASIS_FRONT_DOOR_ENTRY_DOOR_ERROR_DID_NOT_REACH_TARGET_POSITION | 426 | InboxDoor | FrontDoor |  | AssistanceNeeded | AssistanceNeeded | TRUE | An inbox door move has failed. Please try again and call service if it persists | TRUE | Failed to open/close the door, push the doors slowly in a couple of inches. Press the open/close door button again. If the problem persists call service. | Assistance Needed | Y
ERROR_CHASIS_FRONT_DOOR_EXIT_DOOR_ERROR_DID_NOT_REACH_TARGET_POSITION | ERROR_CHASIS_FRONT_DOOR_EXIT_DOOR_ERROR_DID_NOT_REACH_TARGET_POSITION | 427 | OutboxDoor | FrontDoor |  | Warning | AssistanceNeeded | FALSE |  | TRUE | Failed to open/close the door, push the doors slowly in a couple of inches. Press the open/close door button again. If the problem persists call service. | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_BAD_TAG_TYPE | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_BAD_TAG_TYPE | 428 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | AssistanceNeeded | AssistanceNeeded | TRUE | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | TRUE | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_CORRUPT_TAG | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_CORRUPT_TAG | 429 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | AssistanceNeeded | AssistanceNeeded | TRUE | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | TRUE | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_BUS_ERROR | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_BUS_ERROR | 430 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, bus error | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TIMEOUT | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TIMEOUT | 431 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, timeout | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_CORRUPT_RO_SECTION | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_CORRUPT_RO_SECTION | 432 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | AssistanceNeeded | Warning | TRUE | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | TRUE | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_CORRUPT_RW_SECTION | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_CORRUPT_RW_SECTION | 433 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | AssistanceNeeded | Warning | TRUE | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | TRUE | An error has occurred while scanning the plate. Please try a different plate. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_UNKNOWN_ERROR | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_UNKNOWN_ERROR | 449 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, unknown error | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_CHANGED_BEFORE_WRITE | 435 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, tag changed before write | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_WRITE_DATA_FAILED | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_WRITE_DATA_FAILED | 436 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, write data failed | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_WRITE_DATA_BUFFER_CORRUPT | 437 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, write data buffer corrupt | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_RESETTING_MODULE_POWER | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_RESETTING_MODULE_POWER | 438 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, resetting module power | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_POWER_FAULT | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_POWER_FAULT | 439 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, power fault | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_I2C_RFID_BUS_DOWN | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_I2C_RFID_BUS_DOWN | 440 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, I2C RFID bus down | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_I2C_FRONT_LED_BUS_DOWN | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_I2C_FRONT_LED_BUS_DOWN | 441 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, I2C front LED bus down | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_RESETTING_I2C_PERIPHERAL | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_RESETTING_I2C_PERIPHERAL | 442 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE | Chassis front door RFID error, resetting I2C peripheral | TRUE | Chassis front door RFID error, resetting I2C peripheral | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR | 443 | Instrument | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | FALSE |  | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG | 444 | Instrument | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | FALSE |  | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT | 445 | Instrument | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | FALSE |  | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT | 446 | Instrument | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | FALSE |  | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG | 447 | Instrument | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | FALSE |  | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_MODULE_IS_HALTED | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_MODULE_IS_HALTED | 448 | Instrument | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | FALSE |  | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_UNKNOWN_ERROR | ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_UNKNOWN_ERROR | 449 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door RFID error, unknown error | Warning | N
ERROR_CHASIS_POWER_CAN_DIST_INTERLOCK_ACTIVE | ERROR_CHASIS_POWER_CAN_DIST_INTERLOCK_ACTIVE | 450 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE |  | TRUE | The instrument front panel has been opened. All runs have been aborted. Please close the front panel and power cycle the instrument. | Service Call | Y
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_P1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_P1 | 451 | RfidProbesOil1 | RfidProbesOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID invalid firmware version P1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_P2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_P2 | 452 | RfidProbesOil2 | RfidProbesOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID invalid firmware version P2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_E1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_E1 | 453 | RfidEvagreenOil1 | RfidEvagreenOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID invalid firmware version E1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_E2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_E2 | 454 | RfidEvagreenOil2 | RfidEvagreenOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID invalid firmware version E2 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_D1 | ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_D1 | 455 | RfidDrOil1 | RfidDrOil1,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID invalid firmware version D1 | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_D2 | ERROR_CHASIS_ENTRY_HOTEL_RFID_INVALID_FW_VERSION_D2 | 456 | RfidDrOil2 | RfidDrOil2,Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID invalid firmware version D2 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_INVALID_FW_VERSION_W1 | ERROR_CHASIS_EXIT_HOTEL_RFID_INVALID_FW_VERSION_W1 | 457 | RfidWaste1 | RfidWaste1,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID invalid firmware version W1 | Warning | N
ERROR_CHASIS_EXIT_HOTEL_RFID_INVALID_FW_VERSION_W2 | ERROR_CHASIS_EXIT_HOTEL_RFID_INVALID_FW_VERSION_W2 | 458 | RfidWaste2 | RfidWaste2,Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID invalid firmware version W2 | Warning | N
ERROR_CHASIS_FRONT_DOOR_RFID_INVALID_FW_VERSION | ERROR_CHASIS_FRONT_DOOR_RFID_INVALID_FW_VERSION | 459 | FrontDoorRfid | FrontDoorRfid,FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel RFID invalid firmware version | Warning | N
ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_RFID_INVALID_FW_VERSION | ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_RFID_INVALID_FW_VERSION | 460 | Instrument | BottleAndFanController,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller RFID invalid firmware version | Warning | N
ERROR_MISSING_NODE_MASTER_PLACEHOLDER | ERROR_MISSING_NODE_MASTER_PLACEHOLDER | 461 | Instrument | Instrument |  | Warning | Warning | FALSE |  | TRUE | Missing node master placeholder | Warning | N
ERROR_MISSING_NODE_INSTRUMENT_MASTER | ERROR_MISSING_NODE_INSTRUMENT_MASTER | 462 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node instrument master | Service Call | Y
ERROR_MISSING_NODE_AXIS_X | ERROR_MISSING_NODE_AXIS_X | 463 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node axis X | Service Call | Y
ERROR_MISSING_NODE_AXIS_Y | ERROR_MISSING_NODE_AXIS_Y | 464 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node axis Y | Service Call | Y
ERROR_MISSING_NODE_AXIS_Z | ERROR_MISSING_NODE_AXIS_Z | 465 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node axis Z | Service Call | Y
ERROR_MISSING_NODE_PISTON_PUMP_1 | ERROR_MISSING_NODE_PISTON_PUMP_1 | 466 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node piston pump 1 | Service Call | Y
ERROR_MISSING_NODE_PISTON_PUMP_2 | ERROR_MISSING_NODE_PISTON_PUMP_2 | 467 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node piston pump 2 | Service Call | Y
ERROR_MISSING_NODE_DETECTOR_LED_DRIVE | ERROR_MISSING_NODE_DETECTOR_LED_DRIVE | 468 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node detector LED drive | Service Call | Y
ERROR_MISSING_NODE_DR_FLUIDICS_CONTROLLER | ERROR_MISSING_NODE_DR_FLUIDICS_CONTROLLER | 469 | Instrument | DR |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node DR fluidics controller | Service Call | Y
ERROR_MISSING_NODE_BOTTLE_SWITCHING_VALVE_CONTROLLER | ERROR_MISSING_NODE_BOTTLE_SWITCHING_VALVE_CONTROLLER | 470 | Instrument | Bottles |  | Warning | Warning | FALSE |  | TRUE | Missing node bottle switching valve controller | Warning | N
ERROR_MISSING_NODE_ENTRY_HOTEL | ERROR_MISSING_NODE_ENTRY_HOTEL | 471 | Instrument | Inbox |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node entry hotel | Service Call | Y
ERROR_MISSING_NODE_EXIT_HOTEL | ERROR_MISSING_NODE_EXIT_HOTEL | 472 | Instrument | Outbox |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node exit hotel | Service Call | Y
ERROR_MISSING_NODE_POWER_CAN_DIST | ERROR_MISSING_NODE_POWER_CAN_DIST | 473 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | A Communications error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Missing node CAN bus power distributor | Service Call | Y
ERROR_MISSING_NODE_BUS_SENTINEL | ERROR_MISSING_NODE_BUS_SENTINEL | 474 | Instrument | Instrument |  | Warning | Warning | FALSE |  | TRUE | Missing node bus sentinel | Warning | N
ERROR_DR_INSTRUMENT_MASTER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_INSTRUMENT_MASTER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 475 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR instrument master did not respond to heartbeat request | Warning | N
ERROR_DR_AXIS_X_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_AXIS_X_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 476 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR axis X did not respond to heartbeat request | Warning | N
ERROR_DR_AXIS_Y_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_AXIS_Y_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 477 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR axis Y did not respond to heartbeat request | Warning | N
ERROR_DR_AXIS_Z_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_AXIS_Z_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 478 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR axis Z did not respond to heartbeat request | Warning | N
ERROR_DR_PISTON_PUMP_1_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_PISTON_PUMP_1_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 479 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR piston pump 1 did not respond to heartbeat request | Warning | N
ERROR_DR_PISTON_PUMP_2_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_PISTON_PUMP_2_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 480 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR piston pump 2 did not respond to heartbeat request | Warning | N
ERROR_DR_DETECTOR_LED_DRIVE_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_DETECTOR_LED_DRIVE_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 481 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR detector LED drive did not respond to heartbeat request | Warning | N
ERROR_DG_FLUIDICS_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DG_FLUIDICS_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 482 | DG | DG |  | Warning | Warning | FALSE |  | TRUE | DG fluidics controller did not respond to heartbeat request | Warning | N
ERROR_DR_DETECTOR_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_DETECTOR_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 483 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR detector did not respond to heartbeat request | Warning | N
ERROR_DR_FLUIDICS_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_FLUIDICS_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 484 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR fluidics controller did not respond to heartbeat request | Warning | N
ERROR_DG_LOCAL_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DG_LOCAL_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 485 | DG | DG |  | Warning | Warning | FALSE |  | TRUE | DG local controller did not respond to heartbeat request | Warning | N
ERROR_TC_LOCAL_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_TC_LOCAL_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 486 | TC | TC |  | Warning | Warning | FALSE |  | TRUE | TC local controller did not respond to heartbeat request | Warning | N
ERROR_CHASIS_HANDLER_X_AXIS_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_CHASIS_HANDLER_X_AXIS_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 487 | PlateHandler | PlateHandler |  | Warning | Warning | FALSE |  | TRUE | Chassis handler X axis did not respond to heartbeat request | Warning | N
ERROR_CHASIS_HANDLER_Y_AXIS_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_CHASIS_HANDLER_Y_AXIS_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 488 | PlateHandler | PlateHandler |  | Warning | Warning | FALSE |  | TRUE | Chassis handler Y axis did not respond to heartbeat request | Warning | N
ERROR_CHASIS_HANDLER_Z_AXIS_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_CHASIS_HANDLER_Z_AXIS_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 489 | PlateHandler | PlateHandler |  | Warning | Warning | FALSE |  | TRUE | Chassis handler Z axis did not respond to heartbeat request | Warning | N
ERROR_CHASIS_HANDLER_UNABLE_TO_CHECK_TC_FOR_PLATE | ERROR_CHASIS_HANDLER_UNABLE_TO_CHECK_TC_FOR_PLATE | 490 | PlateHandler | PlateHandler,TC |  | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_CHASIS_HANDLER_UNABLE_TO_REMOVE_PLATE_FROM_DG | ERROR_CHASIS_HANDLER_UNABLE_TO_REMOVE_PLATE_FROM_DG | 491 | PlateHandler | PlateHandler,DG |  | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred while trying to remove a plate from the Droplet Generator. Please turn off the instrument power, open the front panel and remove the plate manually. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred while trying to remove a plate from the Droplet Generator. Please turn off the instrument power, open the front panel and remove the plate manually. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_CHASIS_HANDLER_UNABLE_TO_REMOVE_PLATE_FROM_TC | ERROR_CHASIS_HANDLER_UNABLE_TO_REMOVE_PLATE_FROM_TC | 492 | PlateHandler | PlateHandler,TC |  | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred while trying to remove a plate from the Thermal Cycler. Please turn off the instrument power, open the front panel and remove the plate manually. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred while trying to remove a plate from the Thermal Cycler. Please turn off the instrument power, open the front panel and remove the plate manually. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_CHASIS_HANDLER_UNABLE_TO_REMOVE_PLATE_FROM_DR | ERROR_CHASIS_HANDLER_UNABLE_TO_REMOVE_PLATE_FROM_DR | 493 | PlateHandler | PlateHandler,DR |  | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred while trying to remove a plate from the Droplet Reader. Please turn off the instrument power, open the front panel and remove the plate manually. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred while trying to remove a plate from the Droplet Reader. Please turn off the instrument power, open the front panel and remove the plate manually. If the issue persists, please call Bio-Rad support. | Service Call | Y
ERROR_TC_EJECTOR_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_TC_EJECTOR_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 494 | TC | TC |  | Warning | Warning | FALSE |  | TRUE | TC ejector did not respond to heartbeat request | Warning | N
ERROR_CHASIS_ENTRY_HOTEL_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_CHASIS_ENTRY_HOTEL_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 495 | Inbox | Inbox |  | Warning | Warning | FALSE |  | TRUE | Chassis entry hotel did not respond to heartbeat request | Warning | N
ERROR_DR_MASTER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_DR_MASTER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 496 | DR | DR |  | Warning | Warning | FALSE |  | TRUE | DR master did not respond to heartbeat request | Warning | N
ERROR_CHASIS_POWER_CAN_DIST_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_CHASIS_POWER_CAN_DIST_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 497 | Instrument | Instrument |  | Warning | Warning | FALSE |  | TRUE | Chassis CAN bus power distributor did not respond to heartbeat request | Warning | N
ERROR_CHASIS_EXIT_HOTEL_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_CHASIS_EXIT_HOTEL_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 498 | Outbox | Outbox |  | Warning | Warning | FALSE |  | TRUE | Chassis exit hotel did not respond to heartbeat request | Warning | N
ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 499 | Instrument | BottleAndFanController,Bottles |  | Warning | Warning | FALSE |  | TRUE | Chassis bottle and fan controller did not respond to heartbeat request | Warning | N
ERROR_CHASIS_FRONT_DOOR_BOARD_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | ERROR_CHASIS_FRONT_DOOR_BOARD_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 500 | FrontDoor | FrontDoor |  | Warning | Warning | FALSE |  | TRUE | Chassis front door board did not respond to heartbeat request | Warning | N
BUS_SENTINAL_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | BUS_SENTINAL_DID_NOT_RESPOND_TO_HEARTBEAT_REQUEST | 501 | Instrument | Instrument |  | Warning | Warning | FALSE |  | TRUE | Bus sentinel did not respond to heartbeat request | Warning | N
ERROR_DR_INSTRUMENT_MASTER_ASSERT_DETECTED | ERROR_DR_INSTRUMENT_MASTER_ASSERT_DETECTED | 502 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DR_AXIS_X_ASSERT_DETECTED | ERROR_DR_AXIS_X_ASSERT_DETECTED | 503 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DR_AXIS_Y_ASSERT_DETECTED | ERROR_DR_AXIS_Y_ASSERT_DETECTED | 504 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DR_AXIS_Z_ASSERT_DETECTED | ERROR_DR_AXIS_Z_ASSERT_DETECTED | 505 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DR_PISTON_PUMP_1_ASSERT_DETECTED | ERROR_DR_PISTON_PUMP_1_ASSERT_DETECTED | 506 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DR_PISTON_PUMP_2_ASSERT_DETECTED | ERROR_DR_PISTON_PUMP_2_ASSERT_DETECTED | 507 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DR_DETECTOR_LED_DRIVE_ASSERT_DETECTED | ERROR_DR_DETECTOR_LED_DRIVE_ASSERT_DETECTED | 508 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DG_FLUIDICS_CONTROLLER_ASSERT_DETECTED | ERROR_DG_FLUIDICS_CONTROLLER_ASSERT_DETECTED | 509 | DG | DG |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DR_DETECTOR_ASSERT_DETECTED | ERROR_DR_DETECTOR_ASSERT_DETECTED | 510 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DR_FLUIDICS_CONTROLLER_ASSERT_DETECTED | ERROR_DR_FLUIDICS_CONTROLLER_ASSERT_DETECTED | 511 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DG_LOCAL_CONTROLLER_ASSERT_DETECTED | ERROR_DG_LOCAL_CONTROLLER_ASSERT_DETECTED | 512 | DG | DG |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_TC_LOCAL_CONTROLLER_ASSERT_DETECTED | ERROR_TC_LOCAL_CONTROLLER_ASSERT_DETECTED | 513 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_HANDLER_X_AXIS_ASSERT_DETECTED | ERROR_CHASIS_HANDLER_X_AXIS_ASSERT_DETECTED | 514 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_HANDLER_Y_AXIS_ASSERT_DETECTED | ERROR_CHASIS_HANDLER_Y_AXIS_ASSERT_DETECTED | 515 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_HANDLER_Z_AXIS_ASSERT_DETECTED | ERROR_CHASIS_HANDLER_Z_AXIS_ASSERT_DETECTED | 516 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_TC_EJECTOR_ASSERT_DETECTED | ERROR_TC_EJECTOR_ASSERT_DETECTED | 517 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_ENTRY_HOTEL_ASSERT_DETECTED | ERROR_CHASIS_ENTRY_HOTEL_ASSERT_DETECTED | 518 | Inbox | Inbox |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_DR_MASTER_ASSERT_DETECTED | ERROR_DR_MASTER_ASSERT_DETECTED | 519 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_POWER_CAN_DIST_ASSERT_DETECTED | ERROR_CHASIS_POWER_CAN_DIST_ASSERT_DETECTED | 520 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_EXIT_HOTEL_ASSERT_DETECTED | ERROR_CHASIS_EXIT_HOTEL_ASSERT_DETECTED | 521 | Outbox | Outbox |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_ASSERT_DETECTED | ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_ASSERT_DETECTED | 522 | Instrument | BottleAndFanController,Bottles |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
ERROR_CHASIS_FRONT_DOOR_BOARD_ASSERT_DETECTED | ERROR_CHASIS_FRONT_DOOR_BOARD_ASSERT_DETECTED | 523 | FrontDoor | FrontDoor |  | ServiceCall | ServiceCall | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | Y
BUS_SENTINAL_ASSERT_DETECTED | BUS_SENTINAL_ASSERT_DETECTED | 524 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call | N
ERROR_DR_INSTRUMENT_MASTER_BOOTLOADER_RUNNING | ERROR_DR_INSTRUMENT_MASTER_BOOTLOADER_RUNNING | 525 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DR_AXIS_X_BOOTLOADER_RUNNING | ERROR_DR_AXIS_X_BOOTLOADER_RUNNING | 526 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DR_AXIS_Y_BOOTLOADER_RUNNING | ERROR_DR_AXIS_Y_BOOTLOADER_RUNNING | 527 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DR_AXIS_Z_BOOTLOADER_RUNNING | ERROR_DR_AXIS_Z_BOOTLOADER_RUNNING | 528 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DR_PISTON_PUMP_1_BOOTLOADER_RUNNING | ERROR_DR_PISTON_PUMP_1_BOOTLOADER_RUNNING | 529 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DR_PISTON_PUMP_2_BOOTLOADER_RUNNING | ERROR_DR_PISTON_PUMP_2_BOOTLOADER_RUNNING | 530 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DR_DETECTOR_LED_DRIVE_BOOTLOADER_RUNNING | ERROR_DR_DETECTOR_LED_DRIVE_BOOTLOADER_RUNNING | 531 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DG_FLUIDICS_CONTROLLER_BOOTLOADER_RUNNING | ERROR_DG_FLUIDICS_CONTROLLER_BOOTLOADER_RUNNING | 532 | DG | DG |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Generator error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DR_DETECTOR_BOOTLOADER_RUNNING | ERROR_DR_DETECTOR_BOOTLOADER_RUNNING | 533 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DR_FLUIDICS_CONTROLLER_BOOTLOADER_RUNNING | ERROR_DR_FLUIDICS_CONTROLLER_BOOTLOADER_RUNNING | 534 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DG_LOCAL_CONTROLLER_BOOTLOADER_RUNNING | ERROR_DG_LOCAL_CONTROLLER_BOOTLOADER_RUNNING | 535 | DG | DG |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Generator error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_TC_LOCAL_CONTROLLER_BOOTLOADER_RUNNING | ERROR_TC_LOCAL_CONTROLLER_BOOTLOADER_RUNNING | 536 | TC | TC |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_HANDLER_X_AXIS_BOOTLOADER_RUNNING | ERROR_CHASIS_HANDLER_X_AXIS_BOOTLOADER_RUNNING | 537 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_HANDLER_Y_AXIS_BOOTLOADER_RUNNING | ERROR_CHASIS_HANDLER_Y_AXIS_BOOTLOADER_RUNNING | 538 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_HANDLER_Z_AXIS_BOOTLOADER_RUNNING | ERROR_CHASIS_HANDLER_Z_AXIS_BOOTLOADER_RUNNING | 539 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_TC_EJECTOR_BOOTLOADER_RUNNING | ERROR_TC_EJECTOR_BOOTLOADER_RUNNING | 540 | TC | TC |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_ENTRY_HOTEL_BOOTLOADER_RUNNING | ERROR_CHASIS_ENTRY_HOTEL_BOOTLOADER_RUNNING | 541 | Inbox | Inbox |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call
ERROR_DR_MASTER_BOOTLOADER_RUNNING | ERROR_DR_MASTER_BOOTLOADER_RUNNING | 542 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_POWER_CAN_DIST_BOOTLOADER_RUNNING | ERROR_CHASIS_POWER_CAN_DIST_BOOTLOADER_RUNNING | 543 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_EXIT_HOTEL_BOOTLOADER_RUNNING | ERROR_CHASIS_EXIT_HOTEL_BOOTLOADER_RUNNING | 544 | Outbox | Outbox |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call
ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_BOOTLOADER_RUNNING | ERROR_CHASIS_BOTTLE_AND_FAN_CONTROLLER_BOOTLOADER_RUNNING | 545 | BottleAndFanController | BottleAndFanController,Bottles |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call
ERROR_CHASIS_FRONT_DOOR_BOARD_BOOTLOADER_RUNNING | ERROR_CHASIS_FRONT_DOOR_BOARD_BOOTLOADER_RUNNING | 546 | FrontDoor | FrontDoor |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Control System error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. Please call Bio-Rad support to resolve the issue. | Service Call
BUS_SENTINAL_BOOTLOADER_RUNNING | BUS_SENTINAL_BOOTLOADER_RUNNING | 547 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
HANDLER_LOCATION_CONFIG_NOT_VALID | HANDLER_LOCATION_CONFIG_NOT_VALID | 548 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
HANDLER_INVALID_AXIS | HANDLER_INVALID_AXIS | 549 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
HANDLER_INVALID_LOCATION | HANDLER_INVALID_LOCATION | 550 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
HANDLER_SAVE_LOCATION_FAILED | HANDLER_SAVE_LOCATION_FAILED | 551 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
HANDLER_PLATE_DETECTED_BEFORE_GET_PLATE | HANDLER_PLATE_DETECTED_BEFORE_GET_PLATE | 552 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
HANDLER_BUSY | HANDLER_BUSY | 553 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
HANDLER_TIMEOUT | HANDLER_TIMEOUT | 554 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
HANDLER_PLATE_DETECTED_AFTER_DROP_OFF | HANDLER_PLATE_DETECTED_AFTER_DROP_OFF | 555 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_MW_ASSERTED | ERROR_MW_ASSERTED | 556 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_FRONT_DOOR_ENTRY_NOT_HOMED | ERROR_CHASIS_FRONT_DOOR_ENTRY_NOT_HOMED | 557 | InboxDoor | FrontDoor |  |  | AssistanceNeeded | FALSE |  | TRUE | Entry door motor not homed
ERROR_CHASIS_FRONT_DOOR_EXIT_NOT_HOMED | ERROR_CHASIS_FRONT_DOOR_EXIT_NOT_HOMED | 558 | OutboxDoor | FrontDoor |  |  | AssistanceNeeded | FALSE |  | TRUE | Exit door motor not homed
ERROR_DR_INSTRUMENT_MASTER_CONNECTION_LOST | ERROR_DR_INSTRUMENT_MASTER_CONNECTION_LOST | 559 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
DR_NO_PLATE | DR_NO_PLATE | 560 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_VERSIONS_DO_NOT_MATCH | ERROR_VERSIONS_DO_NOT_MATCH | 561 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DR_BULK_TRANSFER_PATH_LENGTH | ERROR_DR_BULK_TRANSFER_PATH_LENGTH | 562 | DR | DR |  | AssistanceNeeded | ServiceCall | TRUE |  | TRUE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
ERROR_DR_BULK_TRANSFER_FILE_SIZE | ERROR_DR_BULK_TRANSFER_FILE_SIZE | 563 | DR | DR |  | AssistanceNeeded | ServiceCall | TRUE |  | TRUE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
ERROR_DR_BULK_TRANSFER_FILE_NOT_FOUND | ERROR_DR_BULK_TRANSFER_FILE_NOT_FOUND | 564 | DR | DR |  | AssistanceNeeded | ServiceCall | TRUE |  | TRUE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
ERROR_DR_BULK_TRANSFER_WELL_DATA_FAILED | ERROR_DR_BULK_TRANSFER_WELL_DATA_FAILED | 565 | DR | DR |  | AssistanceNeeded | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
ERROR_DR_BULK_TRANSFER_PLATE_DATA_FAILED | ERROR_DR_BULK_TRANSFER_PLATE_DATA_FAILED | 566 | DR | DR |  | AssistanceNeeded | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Assistance Needed | Y
ERROR_CHASIS_BFC_WASTE_OVERFLOW_LOWER_SENSOR_DETECTED | ERROR_CHASIS_BFC_WASTE_OVERFLOW_LOWER_SENSOR_DETECTED | 567 | BottleAndFanController | BottleAndFanController |  | AssistanceNeeded | AssistanceNeeded | TRUE | Waste detected in the overflow reservoir. Please perform aspirate vent overflow procedure before adding new plates. This procedure can be performed while the instrument is processing plates. If the issue persists, please call Bio-Rad support. | TRUE | Waste detected in the overflow reservoir. Please perform the "Aspirate Vent Overflow" procedure before adding new plates. This procedure can be performed while the instrument is processing plates. If the issue persists, please call Bio-Rad support | Assistance Needed | y
ERROR_PLATE_RECOVERY_DATABASE_CORRUPTED | ERROR_PLATE_RECOVERY_DATABASE_CORRUPTED | 568 | Instrument | Instrument |  | Warning | ServiceCall | FALSE |  | TRUE | Error plate recovery database corrupted | Warning
ERROR_CHASIS_BFC_WASTE_OVERFLOW_TOP_SENSOR_DETECTED | ERROR_CHASIS_BFC_WASTE_OVERFLOW_TOP_SENSOR_DETECTED | 569 | BottleAndFanController | BottleAndFanController |  | ServiceCall | AssistanceNeeded | TRUE |  | TRUE | Immediate attention required! Significant waste detected in the overflow reservoir. To continue processing plates, please perform the "Aspirate Vent Overflow" procedure. Failure to do so will cause the system to enter into an error state. If the issue persists, please call Bio-Rad support. | Service Call | y
ERROR_DR_WELL_ACQUISITION_TIMED_OUT | ERROR_DR_WELL_ACQUISITION_TIMED_OUT | 570 | DR | DR |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_POWER_CAN_DIST_POWER_NOT_GOOD | ERROR_CHASIS_POWER_CAN_DIST_POWER_NOT_GOOD | 571 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_AC_VOLTAGE_FAIL | ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_AC_VOLTAGE_FAIL | 572 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_OVER_TEMP | ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_OVER_TEMP | 573 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_24V | ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_24V | 574 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_ES | ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_ES | 575 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_TC | ERROR_CHASIS_POWER_CAN_DIST_FAULT_SIG_TC | 576 | Instrument | TC |  | ServiceCall | ServiceCall | FALSE |  | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_MISSING_NODE_HANDLER_X_AXIS | ERROR_MISSING_NODE_HANDLER_X_AXIS | 577 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_MISSING_NODE_HANDLER_Y_AXIS | ERROR_MISSING_NODE_HANDLER_Y_AXIS | 578 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_MISSING_NODE_HANDLER_Z_AXIS | ERROR_MISSING_NODE_HANDLER_Z_AXIS | 579 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_MISSING_NODE_TC_EJECTOR | ERROR_MISSING_NODE_TC_EJECTOR | 580 | TC | TC |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DG_DG_LOCAL_CONTROLLER_HEATER_KILL_SWITCH_ACTIVE | ERROR_DG_DG_LOCAL_CONTROLLER_HEATER_KILL_SWITCH_ACTIVE | 581 | DG | DG |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DG_DG_LOCAL_CONTROLLER_HEATER_TIMEOUT | ERROR_DG_DG_LOCAL_CONTROLLER_HEATER_TIMEOUT | 582 | DG | DG |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DG_DG_LOCAL_CONTROLLER_THERMOCOUPLE_ERROR | ERROR_DG_DG_LOCAL_CONTROLLER_THERMOCOUPLE_ERROR | 583 | DG | DG |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DG_DG_LOCAL_CONTROLLER_THERMOCOUPLE_1_ERROR | ERROR_DG_DG_LOCAL_CONTROLLER_THERMOCOUPLE_1_ERROR | 584 | DG | DG |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DG_DG_LOCAL_CONTROLLER_THERMOCOUPLE_2_ERROR | ERROR_DG_DG_LOCAL_CONTROLLER_THERMOCOUPLE_2_ERROR | 585 | DG | DG |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_DG_DG_Master_DG_IN_ERROR_STATE | ERROR_DG_DG_Master_DG_IN_ERROR_STATE | 586 | DG | DG |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_CHASIS_HANDLER_MOTION_FAILED | ERROR_CHASIS_HANDLER_MOTION_FAILED | 587 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
ERROR_TC_TC_LOCAL_CONTROLLER_TC_ERROR | ERROR_TC_TC_LOCAL_CONTROLLER_TC_ERROR | 588 | TC | TC |  | ServiceCall | ServiceCall | FALSE |  | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call
UnknownSubmoduleError | UnknownSubmoduleError | 1499 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
InvalidInstrumentSerialNumbers | InvalidInstrumentSerialNumbers | 1501 | Instrument | Instrument |  | ServiceCall | ServiceCall | TRUE | Invalid instrument serial numbers. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Invalid instrument serial numbers. Power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
SW_DG_UnknownError | SW_DG_UnknownError | 1502 | DG | DG |  | ServiceCall | ServiceCall | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Generator error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
SW_TC_UnknownError | SW_TC_UnknownError | 1503 | TC | TC |  | ServiceCall | ServiceCall | TRUE | A Thermal cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Thermal Cycler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
SW_DR_UnknownError | SW_DR_UnknownError | 1504 | DR | DR |  | ServiceCall | ServiceCall | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Droplet Reader error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
SW_PlateHandler_UnknownError | SW_PlateHandler_UnknownError | 1505 | PlateHandler | PlateHandler |  | ServiceCall | ServiceCall | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | A Plate Handler error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
SW_InboxDoor_UnknownError | SW_InboxDoor_UnknownError | 1506 | InboxDoor | InboxDoor |  | ServiceCall | ServiceCall | TRUE | Power cycle the instrument, push the doors slowly in a couple of inches and power on the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Power cycle the instrument, push the doors slowly in a couple of inches and power on the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
SW_OutboxDoor_UnknownError | SW_OutboxDoor_UnknownError | 1507 | OutboxDoor | OutboxDoor |  | ServiceCall | ServiceCall | TRUE | Power cycle the instrument, push the doors slowly in a couple of inches and power on the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Power cycle the instrument, push the doors slowly in a couple of inches and power on the instrument and try again. If the issue persists, please call Bio-Rad support. | Service Call | Y
SW_Instrument_UnknownError | SW_Instrument_UnknownError | 1508 | Instrument | Instrument |  | Yes | ServiceCall | FALSE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | An instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Yes | Yes
SW_Instrument_Offline | SW_Instrument_Offline | 1509 | Instrument | Instrument |  | Yes | ServiceCall | FALSE | Instrument connection has been lost.Please power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | TRUE | Instrument connection has been lost. Please power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | Yes | Yes
SW_InvalidFwVersion | SW_InvalidFwVersion | 1510 | Instrument | Instrument |  | ServiceCall | ServiceCall | FALSE |  | TRUE | The controller firmware version is not compatible with this QX One software version. Please check compatibility matrix and update the QX One software and firmware accordingly. | Service Call
SW_TimeoutWaitingForResponseFromFW | SW_TimeoutWaitingForResponseFromFW | 1511 | Instrument | Instrument |  | ? | ServiceCall | FALSE |  | TRUE | An Instrument error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support. | ?
SoftwareError | SoftwareError | 1999 | Instrument | Instrument |  | Yes | Warning | FALSE | Software Error, exception caught. Some data might be usable.View the Event log for recent software exceptions and report to Bio-Rad support. | TRUE | Software Error, exception caught. Some data might be usable. View the Event log for recent software exceptions and report to Bio-Rad support. | Yes | Yes
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_IS_HALTED_P1 |  | Instrument | RfidProbesOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_IS_HALTED_P2 |  | Instrument | RfidProbesOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_IS_HALTED_E1 |  | Instrument | RfidEvagreenOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_IS_HALTED_E2 |  | Instrument | RfidEvagreenOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_IS_HALTED_D1 |  | Instrument | RfidDrOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_MODULE_IS_HALTED_D2 |  | Instrument | RfidDrOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR_P1 |  | Instrument | RfidProbesOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR_P2 |  | Instrument | RfidProbesOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR_E1 |  | Instrument | RfidEvagreenOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR_E2 |  | Instrument | RfidEvagreenOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR_D1 |  | Instrument | RfidDrOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR_D2 |  | Instrument | RfidDrOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG_P1 |  | Instrument | RfidProbesOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG_P2 |  | Instrument | RfidProbesOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG_E1 |  | Instrument | RfidEvagreenOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG_E2 |  | Instrument | RfidEvagreenOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG_D1 |  | Instrument | RfidDrOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG_D2 |  | Instrument | RfidDrOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT_P1 |  | Instrument | RfidProbesOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT_P2 |  | Instrument | RfidProbesOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT_E1 |  | Instrument | RfidEvagreenOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT_E2 |  | Instrument | RfidEvagreenOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT_D1 |  | Instrument | RfidDrOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT_D2 |  | Instrument | RfidDrOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT_P1 |  | Instrument | RfidProbesOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT_P2 |  | Instrument | RfidProbesOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT_E1 |  | Instrument | RfidEvagreenOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT_E2 |  | Instrument | RfidEvagreenOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT_D1 |  | Instrument | RfidDrOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT_D2 |  | Instrument | RfidDrOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG_P1 |  | Instrument | RfidProbesOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG_P2 |  | Instrument | RfidProbesOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG_E1 |  | Instrument | RfidEvagreenOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG_E2 |  | Instrument | RfidEvagreenOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG_D1 |  | Instrument | RfidDrOil1,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_ENTRY_HOTEL_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG_D2 |  | Instrument | RfidDrOil2,Inbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_MODULE_IS_HALTED_W1 |  | Instrument | RfidWaste1,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_MODULE_IS_HALTED_W2 |  | Instrument | RfidWaste2,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR_W1 |  | Instrument | RfidWaste1,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR_W2 |  | Instrument | RfidWaste2,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG_W1 |  | Instrument | RfidWaste1,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG_W2 |  | Instrument | RfidWaste2,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT_W1 |  | Instrument | RfidWaste1,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT_W2 |  | Instrument | RfidWaste2,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT_W1 |  | Instrument | RfidWaste1,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT_W2 |  | Instrument | RfidWaste2,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG_W1 |  | Instrument | RfidWaste1,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_EXIT_HOTEL_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG_W2 |  | Instrument | RfidWaste2,Outbox |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_BFC_RFID_ERROR_MODULE_IS_HUNG |  | Instrument | BottleAndFanController |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_BFC_RFID_ERROR_MODULE_BUS_DOWN |  | BottleAndFanController | BottleAndFanController |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller RFID error, module bus down
| ERROR_CHASIS_BFC_RFID_ERROR_RESETTING_MODULE_POWER |  | BottleAndFanController | BottleAndFanController |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller RFID error, resetting module power
| ERROR_CHASIS_BFC_RFID_ERROR_RESETTING_I2C_PERIPHERAL |  | BottleAndFanController | BottleAndFanController |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller RFID error, resseting I2C peripheral
| ERROR_CHASIS_BFC_RFID_ERROR_MODULE_FAULT |  | Instrument | BottleAndFanController |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_BFC_RFID_ERROR_TAG_ERROR_RFID_MODULE_RETURNED_ERROR |  | Instrument | BottleAndFanController |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_BFC_RFID_ERROR_TAG_ERROR_UUID_SIZE_TOO_BIG |  | Instrument | BottleAndFanController |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_BFC_RFID_ERROR_TAG_ERROR_RFID_MODULE_COMMAND_CORRUPT |  | Instrument | BottleAndFanController |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_BFC_RFID_ERROR_TAG_ERROR_RFID_MODULE_TIMED_OUT |  | Instrument | BottleAndFanController |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_BFC_RFID_ERROR_TAG_ERROR_RFID_COMMAND_TOO_LONG |  | Instrument | BottleAndFanController |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_BFC_RFID_ERROR_MODULE_IS_HALTED |  | Instrument | BottleAndFanController |  |  | Warning |  |  | FALSE
| ERROR_CHASIS_BFC_ROTATOR_LOST_COMS |  | BottleAndFanController | BottleAndFanController |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller rotator lost communication
| ERROR_CHASIS_BFC_ROTATOR_TIMEOUT |  | BottleAndFanController | BottleAndFanController |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller rotator timeout
| ERROR_CHASIS_BFC_OIL_BOTTLE_REMOVED_WHILE_IN_USE_P1 |  | BottleAndFanController | BottleAndFanController,Bottles |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller oil bottle removed while in use P1
| ERROR_CHASIS_BFC_OIL_BOTTLE_REMOVED_WHILE_IN_USE_P2 |  | BottleAndFanController | BottleAndFanController,Bottles |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller oil bottle removed while in use P2
| ERROR_CHASIS_BFC_OIL_BOTTLE_REMOVED_WHILE_IN_USE_E1 |  | BottleAndFanController | BottleAndFanController,Bottles |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller oil bottle removed while in use E1
| ERROR_CHASIS_BFC_OIL_BOTTLE_REMOVED_WHILE_IN_USE_E2 |  | BottleAndFanController | BottleAndFanController,Bottles |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller oil bottle removed while in use E2
| ERROR_CHASIS_BFC_OIL_BOTTLE_REMOVED_WHILE_IN_USE_D1 |  | BottleAndFanController | BottleAndFanController,Bottles |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller oil bottle removed while in use D1
| ERROR_CHASIS_BFC_OIL_BOTTLE_REMOVED_WHILE_IN_USE_D2 |  | BottleAndFanController | BottleAndFanController,Bottles |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller oil bottle removed while in use D2
| ERROR_CHASIS_BFC_WASTE_BOTTLE_REMOVED_WHILE_IN_USE_W1 |  | BottleAndFanController | BottleAndFanController,Bottles |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller waste bottle removed while in use W1
| ERROR_CHASIS_BFC_WASTE_BOTTLE_REMOVED_WHILE_IN_USE_W2 |  | BottleAndFanController | BottleAndFanController,Bottles |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller waste bottle removed while in use W2
| ERROR_CHASIS_BFC_MESSAGE_DROP_DETECTED |  | BottleAndFanController | BottleAndFanController |  |  | Warning |  |  | TRUE | Chassis bottle and fan controller message drop detected
| ERROR_CHASIS_FRONT_DOOR_RFID_ERROR_MODULE_IS_HUNG |  | Instrument | FrontDoorRfid,FrontDoor |  |  | Warning |  |  | FALSE
| SW_FrontDoorRfid_UnknownError |  | FrontDoorRfid | FrontDoorRfid |  |  | ServiceCall |  |  | TRUE | A Front Door RFID error has occurred. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
| SW_Slot_UnknownError_Slot1 |  | Slot | Slot |  |  | ServiceCall |  |  | TRUE | An error has occurred with Slot 1. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
| SW_Slot_UnknownError_Slot2 |  | Slot | Slot |  |  | ServiceCall |  |  | TRUE | An error has occurred with Slot 2. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
| SW_Slot_UnknownError_Slot3 |  | Slot | Slot |  |  | ServiceCall |  |  | TRUE | An error has occurred with Slot 3. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
| SW_Slot_UnknownError_Slot4 |  | Slot | Slot |  |  | ServiceCall |  |  | TRUE | An error has occurred with Slot 4. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
| SW_Slot_UnknownError_Slot5 |  | Slot | Slot |  |  | ServiceCall |  |  | TRUE | An error has occurred with Slot 5. Please allow the current runs to complete, if possible. Then power cycle the instrument and try again. If the issue persists, please call Bio-Rad support.
Note on N1: Marked as orange for a draft response. Go green when it has been checked Vs FW inputs in columns C and D. -Deleted user
Note on A18: At what point does this get generated? does it stop the handler? -Deleted user

## Summary
Summary | Hello | Hello
Original Error Codes Repeated
JSON Error Codes Repeated
534
|  |  | CW | CW

## Variables
text#text#text |  | 7 |  |  | 1 | 3 | 0.33333
text text text |  | 2
Service Call |  | 5 | 2
ServiceCall |  | -3 | 3
|  | 8 | -1 | 0
| IBM | 2
| Google | 8
Apple
Apple
Google
TikTok
| Adobe | [6 empty] | Service Call | ServiceCall | ServiceCall
[8 empty] | Assistance Needed | AssistanceNeeded | ServiceCall
2
'InboxSlot': lambda s: ALL(s, ['ENTRY', 'HOTEL']) and ANY(s, ['SLOT1', 'SLOT2', 'SLOT3', 'SLOT4', 'SLOT5']), | InboxSlot
'OutboxSlot': lambda s: ALL(s, ['EXIT', 'HOTEL']) and ANY(s, ['SLOT1', 'SLOT2', 'SLOT3', 'SLOT4', 'SLOT5']), | OutboxSlot
'RfidProbesOil1': lambda s: ALL(s, ['RFID', 'P1']), | RfidProbesOil1
'RfidProbesOil2': lambda s: ALL(s, ['RFID', 'P2']), | RfidProbesOil2
'RfidEvagreenOil1': lambda s: ALL(s, ['RFID', 'E1']), | RfidEvagreenOil1
'RfidEvagreenOil2': lambda s: ALL(s, ['RFID', 'E2']), | RfidEvagreenOil2
'RfidWaste1': lambda s: ALL(s, ['RFID', 'W1']), | RfidWaste1
'RfidWaste2': lambda s: ALL(s, ['RFID', 'W2']), | RfidWaste2
'RfidDrOil1': lambda s: ALL(s, ['RFID', 'D1']), | RfidDrOil1
'RfidDrOil2': lambda s: ALL(s, ['RFID', 'D2']), | RfidDrOil2
'BottleAndFanController': lambda s: 'BFC' in s or ALL(s, ['BOTTLE', 'AND', 'FAN', 'CONTROLLER']), | BottleAndFanController
'FrontDoorRfid': lambda s: ALL(s, ['FRONT', 'DOOR', 'RFID']) or 'FrontDoorRfid' in s, | FrontDoorRfid
'Inbox': lambda s: ALL(s, ['ENTRY', 'HOTEL']), | Inbox
'Outbox': lambda s: ALL(s, ['EXIT', 'HOTEL']), | Outbox
'FrontDoor': lambda s: ALL(s, ['FRONT', 'DOOR']), | FrontDoor
'Slot': lambda s: ANY(s, ['SLOT1', 'SLOT2', 'SLOT3', 'SLOT4', 'SLOT5']) or ANY(s, ['Slot1', 'Slot2', 'Slot3', 'Slot4', 'Slot5']), | Slot
'PlateHandler': lambda s: 'HANDLER' in s or 'PlateHandler' in s or 'PH' in s, | PlateHandler
'DG': lambda s: 'DG' in s, | DG
'DR': lambda s: 'DR' in s and 'DROPLET_GENERATION' not in s, | DR
'TC': lambda s: 'TC' in s and 'CATCH' not in s, | TC
'Instrument': lambda s: 'CONTROLLER' in s, | Instrument
'Bottles': lambda s: 'BOTTLE' in s or ALL(s, ['INSUFFICIENT', 'OIL']) or ALL(s, ['INSUFFICIENT', 'WASTE']), | Bottles
'InboxDoor': lambda s: 'InboxDoor' in s, | InboxDoor
'OutboxDoor': lambda s: 'OutboxDoor' in s, | OutboxDoor